## Result Concentration
Header_1 | Measurement Date | Calculation Date | Method | Operator | Wavelength Drift Status | Dilution Factor | Calc_Mode | SampleName | Na 589.592 | Na 588.995 | Na 589.592 R | Na 330.298 R | K 766.491 | K 769.896 | K 766.491 R | K 769.896 R | Mg 285.213 | Mg 279.079 | Mg 279.079 R | Mg 285.213 R | Ca 393.366 | Ca 396.847 | Ca 317.933 | Ca 422.673 | Ar 404.442 | C 193.091 | Y 224.306 | Y 371.030 | Y 324.228 | Al 167.078 | Al 396.152 | As 189.042 | As 193.759 | B 249.773 | B 249.677 | Ba 233.527 | Ba 455.404 | Be 234.861 | Be 313.042 | Be 313.107 | Cd 214.438 | Cd 226.502 | Co 228.616 | Co 238.892 | Cr 267.716 | Cr 283.563 | Cu 324.754 | Cu 224.700 | Cu 327.396 | Fe 238.204 | Fe 259.941 | Li 670.780 | Li 670.780 | Mn 257.611 | Mn 259.373 | Ni 221.648 | Ni 231.604 | Ni 232.003 | Pb 220.353 | Pb 283.305 | Se 196.090 | Se 204.050 | Sr 407.771 | Sr 421.552 | Zn 213.856 | Zn 206.200 | Dilution
Header_2 | [8 empty] | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | ppm | cps | cps | cps | cps | cps | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb | ppb
Header_3 | [8 empty] | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Stdz_Int | Stdz_Int | Stdz_Int | Stdz_Int | Stdz_Int | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc | Conc
Mean | 02/07/2026 13:26:07 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV 1 | 1.01 | 0.986 | 4.25 | 6.12 | 1.05 | 1.04 | <2.31 | <2.79 | 1.05 | 1.10 | 1.17 | 0.248 | 1.01 | 1.00 | 1.08 | 1.14 | 2351158 | 480500 | 271349 | 737653 | 480425 | 1096 | 1019 | 994 | 1018 | 999 | 1030 | 1035 | 1043 | 1010 | 1083 | 1057 | 1063 | 1064 | 1043 | 1033 | 1069 | 1048 | 1027 | 1036 | 1064 | 1048 | 1023 | 1014 | 1014 | 1068 | 1019 | 1045 | 1089 | 1048 | 1044 | 1018 | 1051 | 1021 | 1012 | 1056 | 1017 | 1050 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 1
Mean | 02/07/2026 10:01:21 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | RE_AC_452 | <-321 | <-53.8 | 116 | 120 | 6.43 | >6.13 | 12.4 | 12.9 | 32.5 | 30.5 | 30.0 | 31.1 | >2.60 | 5.25 | >243 | >241 | 2280104 | 54435098 | 259312 | 616111 | 402000 | 3.65 | <-4.91 | 15.1 | 16.5 | 192 | 189 | 240 | 244 | <-0.125 | 0.147 | 0.0488 | <-0.004162 | <-0.126 | 1.61 | 1.44 | 4.61 | 4.74 | 0.377 | 0.672 | <-0.480 | 160 | 153 | 27.4 | 27.4 | 910 | 885 | 11.0 | 5.65 | 5.46 | 1.80 | 64.8 | 23.0 | <-36.6 | 861 | 858 | 4.20 | 2.58 | RE_AC_452 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | RE_AC_452 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | RE_AC_452 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | RE_AC_452 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | RE_AC_452 | 1
Mean | 02/07/2026 09:30:09 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Blank | 0.003040 | 0.000180 | 0.868 | 3.10 | 0.008770 | 0.001591 | <1.41 | <2.05 | 0.001328 | 0.007378 | 0.143 | <-0.907 | 0.000949 | 0.000392 | <-0.001076 | 0.0607 | 2336921 | 363201 | 274683 | 747613 | 486800 | 0.0403 | <-7.80 | 0.395 | 15.7 | 9.95 | 8.80 | <-0.003936 | <-0.0588 | 0.009176 | <-0.0199 | 0.0102 | <-0.0220 | 0.000422 | 0.705 | 0.585 | <-0.0908 | <-0.322 | <-0.0341 | 0.249 | <-0.0280 | 0.0725 | <-6.19 | 0.0253 | 0.001938 | 0.0166 | 0.0432 | 0.0456 | <-0.009371 | <-0.113 | <-0.160 | 61.7 | 6.36 | <-0.532 | 0.0129 | <-61.5 | 0.0768 | <-1.48 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Blank | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Blank | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 09:34:35 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV 1 | 1.04 | 1.02 | 1.33 | 8.25 | 1.06 | 1.04 | <1.80 | <2.24 | 1.04 | 1.09 | 1.19 | 0.260 | 1.03 | 1.01 | 1.06 | 1.11 | 2293652 | 300475 | 270337 | 743777 | 469313 | 1053 | 1035 | 1018 | 998 | 1001 | 1010 | 1029 | 1081 | 1013 | 1044 | 1027 | 1045 | 1036 | 1039 | 1022 | 1035 | 1029 | 1033 | 1041 | 1049 | 1033 | 1027 | 1037 | 1037 | 1054 | 1026 | 1036 | 1056 | 1028 | 1045 | 1021 | 1024 | 1006 | 1036 | 1039 | 1011 | 1040 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 09:39:03 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV_50 | <-38.5 | 1.71 | 46.7 | 50.3 | <-74.3 | <-107 | 48.7 | 49.3 | 52.5 | 53.2 | 50.5 | 52.9 | >2.79 | 5.48 | 50.2 | 49.3 | 2319342 | 324670 | 265490 | 713991 | 465805 | 13.5 | 4.99 | 1.37 | 13.3 | 8.99 | 7.05 | 0.497 | 0.361 | 0.0986 | 0.0392 | 0.0225 | 0.0878 | 0.0318 | 0.455 | 0.454 | 0.0514 | 0.120 | 0.941 | 0.531 | 0.823 | 2.09 | <-4.28 | 0.0667 | 0.0422 | 0.181 | 0.220 | 0.349 | 0.414 | 0.456 | 0.700 | 63.7 | 6.14 | <-3.67 | 0.403 | <-61.1 | 2.31 | 0.868 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV_50 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV_50 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 09:43:32 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | MB | 0.133 | 0.129 | 0.512 | 5.89 | <-0.000509 | <-0.008820 | <0.950 | <1.35 | 0.000539 | <-0.000362 | 0.124 | <-0.908 | <-0.006038 | <-0.006864 | <-0.0235 | 0.0545 | 2342806 | 337619 | 263383 | 409872 | 278774 | 0.231 | <-7.40 | <-0.0700 | 14.7 | 1.59 | 1.76 | <-0.0115 | <-0.108 | 0.001882 | 0.0329 | 0.000704 | 0.0102 | <-0.0395 | 0.583 | 0.328 | <-0.108 | 0.155 | 0.794 | 0.240 | 0.331 | <-0.0682 | <-6.33 | 0.0195 | <-0.002148 | <-0.009455 | 0.0156 | 0.0954 | <-0.0741 | <-0.233 | <-0.263 | 63.2 | 5.25 | <-6.25 | 0.009933 | <-61.5 | <-0.259 | <-1.88 | MB | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MB | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | MB | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MB | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | MB | 1
Mean | 02/07/2026 09:47:59 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | STOCK_10 | >6.33 | >6.59 | 8.27 | 13.3 | 2.99 | 2.96 | <3.84 | <4.21 | 15.6 | 15.8 | 15.7 | 15.8 | >2.91 | 5.66 | 36.5 | 35.4 | 2328374 | 544581 | 274648 | 158510 | 117398 | 0.559 | <-6.20 | 49.1 | 62.3 | 101 | 95.2 | 51.0 | 53.0 | 19.3 | 20.4 | 20.0 | 20.0 | 19.7 | 25.5 | 25.6 | 20.1 | 19.4 | 20.3 | 19.8 | 21.0 | 101 | 94.6 | 0.106 | 0.0827 | 31.0 | 30.6 | 49.4 | 50.2 | 48.3 | 24.4 | 85.8 | 10.6 | <-6.82 | 102 | 46.9 | 50.2 | 48.9 | STOCK_10 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | STOCK_10 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | STOCK_10 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | STOCK_10 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | STOCK_10 | 1
Mean | 02/07/2026 09:52:27 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | LFS_1/10 | <-4.27 | >10.2 | 29.1 | 39.6 | <-318 | <-504 | 96.7 | 96.5 | 22.1 | 21.1 | 20.1 | 20.8 | >2.78 | 5.60 | 10.1 | 9.85 | 2329461 | 575965 | 270536 | 729246 | 472650 | >15036 | >20208 | >8309 | 8138 | 3019 | 3017 | 2005 | 2020 | 1999 | >2105 | 2066 | >2041 | 2020 | 2019 | 2010 | 4030 | 4011 | 3048 | >3029 | >3109 | >29564 | >29418 | 2513 | 2513 | 2021 | 2018 | >5038 | 5067 | 4981 | >10100 | 9535 | >20088 | >20106 | >1918 | 2004 | 2023 | 2046 | LFS_1/10 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | LFS_1/10 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | LFS_1/10 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | LFS_1/10 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | LFS_1/10 | 1
Mean | 02/07/2026 09:56:54 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | IHS_54_1/50 | <-33.4 | 3.74 | 49.2 | 52.9 | 1.71 | 1.66 | <2.24 | <2.84 | 6.69 | 6.57 | 6.54 | 5.95 | >2.61 | 5.28 | 6.29 | 6.18 | 2332618 | 1015212 | 286936 | 588111 | 384445 | 1.69 | <-6.25 | 1.22 | 16.0 | 36.0 | 34.2 | 2.17 | 2.18 | 0.0188 | <-0.009570 | <-0.0186 | 0.0527 | 0.0235 | 0.746 | 0.842 | 0.508 | 0.0213 | <-0.382 | 0.237 | 1.18 | 2.22 | <-4.11 | 1.09 | 1.06 | 9.50 | 9.27 | 0.408 | 0.196 | 0.277 | 0.897 | 65.3 | 8.03 | <-1.99 | 182 | 132 | 3.60 | 2.22 | IHS_54_1/50 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_54_1/50 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | IHS_54_1/50 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_54_1/50 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | IHS_54_1/50 | 1
Mean | 02/07/2026 10:05:48 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | 55429_1554316 | <-396 | <-66.6 | 127 | 132 | 3.60 | 3.48 | <3.85 | <4.37 | 34.7 | 33.1 | 31.7 | 33.1 | >2.83 | 5.70 | 174 | 171 | 2317060 | 21848646 | 247942 | 536965 | 355924 | 1.34 | <-8.22 | 6.15 | 14.4 | 77.9 | 74.5 | 152 | 157 | <-0.006633 | 0.0981 | <-0.0329 | 0.0190 | <-0.0873 | 1.57 | 1.44 | 1.88 | 2.42 | 13.3 | 12.4 | 11.8 | 1.27 | <-5.08 | 3.08 | 3.04 | 2.95 | 2.91 | 4.07 | 0.193 | 0.406 | 1.16 | 64.3 | 13.3 | <-20.1 | 750 | 742 | 25.8 | 24.4 | 55429_1554316 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | 55429_1554316 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | 55429_1554316 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | 55429_1554316 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 55429_1554316 | 1
Mean | 02/07/2026 10:10:15 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_as is | <-6.27 | >9.85 | 35.0 | 38.7 | 1.94 | 1.90 | <2.64 | <3.18 | 11.7 | 11.7 | 11.4 | 11.2 | >2.78 | 5.46 | 48.6 | 47.6 | 2339794 | 1901494 | 266150 | 714258 | 464991 | 3.69 | <-4.67 | 0.835 | 15.0 | 208 | 204 | 39.5 | 40.6 | 0.007618 | 0.0181 | 0.0104 | 0.0373 | 0.0135 | 1.05 | 1.05 | 0.349 | 0.435 | 3.53 | 3.55 | 3.14 | 0.780 | <-5.41 | 2.45 | 2.41 | 0.737 | 0.756 | 1.49 | 1.01 | 1.01 | 0.468 | 64.0 | 8.95 | <-6.69 | 255 | 210 | 595 | 613 | Var_9_as is | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_as is | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_as is | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_as is | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_as is | 1
Mean | 02/07/2026 10:14:42 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_dup | <-4.27 | >10.3 | 33.9 | 40.0 | 1.92 | 1.87 | <2.57 | <3.08 | 11.9 | 11.8 | 11.5 | 11.1 | >2.76 | 5.42 | 47.6 | 46.8 | 2334025 | 1417399 | 268221 | 709181 | 459600 | 3.11 | <-5.24 | 1.36 | 14.4 | 209 | 207 | 39.5 | 39.7 | 0.0105 | 0.0104 | <-0.0176 | 0.0464 | <-0.0193 | 1.03 | 0.998 | 0.402 | 0.296 | 3.68 | 3.36 | 3.30 | 0.387 | <-5.99 | 2.41 | 2.38 | 0.641 | 0.636 | 1.45 | 0.923 | 0.447 | <-0.175 | 62.7 | 6.65 | <-5.03 | 248 | 203 | 611 | 627 | Var_9_dup | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_dup | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_dup | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_dup | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_dup | 1
Mean | 02/07/2026 10:19:10 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_ms_Na | <-335 | <-51.0 | 122 | 128 | 2.13 | 2.05 | <2.56 | <3.15 | 11.1 | 10.5 | 10.1 | 9.74 | >2.78 | 5.45 | 42.6 | 41.7 | 2349088 | 1289308 | 267042 | 704264 | 461581 | 3.25 | <-6.09 | 0.779 | 14.2 | 198 | 195 | 34.9 | 35.0 | 0.008835 | 0.0152 | 0.004045 | 0.0733 | 0.0123 | 1.15 | 1.18 | 0.389 | 0.411 | 3.36 | 3.07 | 3.15 | 1.61 | <-4.76 | 2.41 | 2.38 | 0.643 | 0.701 | 1.72 | 1.37 | 1.55 | 0.212 | 64.5 | 6.92 | <-2.30 | 217 | 169 | 500 | 508 | Var_9_ms_Na | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_Na | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_ms_Na | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_Na | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_ms_Na | 1
Mean | 02/07/2026 10:23:38 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_ms_K | <-6.08 | >9.88 | 34.7 | 38.7 | 5.45 | 5.47 | 6.48 | 6.91 | 12.0 | 11.9 | 11.4 | 11.5 | >2.82 | 5.52 | 48.9 | 47.8 | 2340224 | 625830 | 263721 | 706255 | 460735 | 3.88 | <-5.75 | 0.727 | 14.2 | 213 | 210 | 40.0 | 41.2 | <-0.0214 | 0.0201 | 0.0180 | 0.0669 | 0.0122 | 1.01 | 1.19 | 0.434 | 0.513 | 4.67 | 4.18 | 3.67 | 26.7 | 20.3 | 2.56 | 2.52 | 0.907 | 0.967 | 1.59 | 0.678 | 1.00 | 0.0862 | 64.0 | 7.55 | <-4.60 | 255 | 210 | 618 | 630 | Var_9_ms_K | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_K | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_ms_K | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_K | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_ms_K | 1
Mean | 02/07/2026 10:28:06 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_ms_Mg+Ca | <-2.38 | >10.6 | 30.4 | 36.7 | 1.91 | 1.87 | <2.37 | <2.76 | 42.0 | 42.0 | 39.4 | 41.4 | >2.63 | 5.31 | 142 | 141 | 2308747 | 556005 | 264903 | 713328 | 460301 | 7.58 | <-2.43 | 1.69 | 14.7 | 192 | 188 | 34.3 | 34.8 | <-0.006485 | 0.0194 | 0.003435 | 0.0928 | 0.009027 | 1.10 | 1.20 | 0.559 | 0.403 | 4.37 | 4.05 | 3.62 | 10.8 | 4.46 | 2.35 | 2.31 | 0.909 | 0.929 | 1.46 | 0.908 | 0.667 | 0.824 | 65.9 | 6.77 | <-4.66 | 217 | 169 | 531 | 533 | Var_9_ms_Mg+Ca | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_Mg+Ca | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_ms_Mg+Ca | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_Mg+Ca | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_ms_Mg+Ca | 1
Mean | 02/07/2026 10:32:34 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_ms_met_100 ppb | <-6.59 | >9.76 | 34.8 | 41.2 | 2.11 | 2.07 | <2.53 | <2.84 | 12.1 | 12.1 | 11.5 | 11.6 | >2.84 | 5.56 | 48.8 | 48.1 | 2327949 | 1740153 | 262172 | 710117 | 459460 | 105 | 93.2 | 102 | 116 | 311 | 312 | 138 | 143 | 99.5 | 104 | 102 | 101 | 100 | 99.8 | 99.2 | 101 | 98.5 | 104 | 102 | 105 | 99.9 | 93.1 | 123 | 123 | 101 | 98.0 | 100 | 100 | 98.8 | 97.9 | 164 | 110 | 94.8 | 353 | 316 | 717 | 733 | Var_9_ms_met_100 ppb | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_met_100 ppb | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_ms_met_100 ppb | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_met_100 ppb | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_ms_met_100 ppb | 1
Mean | 02/07/2026 10:37:03 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_ms_met_100 ppb_d | <-6.27 | >9.86 | 34.0 | 38.3 | 2.07 | 2.02 | <2.49 | <2.88 | 12.0 | 11.9 | 11.4 | 11.4 | >2.78 | 5.45 | 48.6 | 47.5 | 2335606 | 848965 | 266823 | 712281 | 464980 | 105 | 94.5 | 104 | 116 | 306 | 306 | 137 | 141 | 98.8 | 103 | 101 | 100 | 99.3 | 98.8 | 98.3 | 99.2 | 97.6 | 103 | 102 | 103 | 101 | 94.8 | 122 | 122 | 101 | 97.9 | 99.8 | 101 | 98.6 | 97.7 | 172 | 108 | 91.7 | 354 | 311 | 706 | 722 | Var_9_ms_met_100 ppb_d | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_met_100 ppb_d | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_ms_met_100 ppb_d | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_met_100 ppb_d | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_ms_met_100 ppb_d | 1
Mean | 02/07/2026 10:41:32 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Var_9_ms_met_200 ppb_B | <-5.01 | >10.1 | 33.9 | 38.0 | 2.16 | 2.12 | <2.58 | <2.89 | 12.0 | 11.9 | 11.5 | 11.2 | >2.78 | 5.45 | 47.9 | 47.0 | 2338322 | 661551 | 266517 | 716234 | 465770 | 206 | 192 | 201 | 213 | 416 | 417 | 235 | 242 | 198 | 208 | 204 | 201 | 199 | 197 | 195 | 199 | 197 | 202 | 200 | 204 | 199 | 191 | 242 | 242 | 201 | 195 | 198 | 201 | 195 | 195 | 265 | 210 | 193 | 444 | 414 | 788 | 803 | Var_9_ms_met_200 ppb_B | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_met_200 ppb_B | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Var_9_ms_met_200 ppb_B | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Var_9_ms_met_200 ppb_B | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Var_9_ms_met_200 ppb_B | 1
Mean | 02/07/2026 10:46:00 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Blank | <-0.000233 | <-0.003327 | 0.334 | 4.12 | 0.004484 | <-0.002172 | <0.802 | <1.15 | 0.001532 | <-0.000119 | 0.112 | <-0.907 | 0.001492 | 0.000866 | 0.000857 | 0.0716 | 2337160 | 430624 | 271977 | 736027 | 476433 | 0.0920 | <-7.81 | 0.666 | 14.3 | 3.97 | 2.15 | 0.0280 | <-0.0748 | 0.0185 | <-0.001685 | 0.0433 | 0.0473 | <-0.0297 | 1.01 | 1.05 | 0.0538 | <-0.151 | 0.232 | 0.273 | <-0.267 | 0.117 | <-6.17 | <-0.000954 | <-0.0319 | <-0.004107 | 0.0317 | <-0.0497 | <-0.0378 | 0.142 | 0.0210 | 64.3 | 4.88 | <-0.0348 | 0.0108 | <-61.5 | 0.0624 | <-1.36 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Blank | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Blank | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 10:50:26 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV 1 | 0.977 | 0.975 | 1.23 | 9.32 | 1.000 | 0.991 | <1.70 | <2.02 | 1.04 | 1.08 | 1.22 | 0.250 | 0.987 | 0.967 | 1.03 | 1.09 | 2340889 | 346758 | 270524 | 714689 | 462854 | 1071 | 977 | 1007 | 1002 | 1000 | 1015 | 1026 | 1027 | 1014 | 1039 | 1018 | 1056 | 1044 | 1038 | 1025 | 1045 | 1029 | 1007 | 1038 | 1014 | 1036 | 1023 | 980 | 980 | 1057 | 1026 | 1047 | 1068 | 1027 | 1045 | 1010 | 1027 | 1005 | 995 | 1006 | 1012 | 1040 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 10:54:54 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV_50 | <-40.9 | 1.09 | 46.1 | 46.0 | <-78.5 | <-104 | 48.7 | 48.9 | 52.6 | 53.5 | 50.1 | 51.9 | >2.79 | 5.47 | 49.7 | 49.0 | 2337362 | 361964 | 265736 | 715463 | 464747 | 13.6 | 3.80 | 0.463 | 16.3 | 7.85 | 6.07 | 0.379 | 0.329 | 0.0313 | 0.0634 | 0.0185 | 0.0508 | <-0.0133 | 0.906 | 0.906 | 0.128 | 0.0688 | 0.919 | 0.183 | <-0.119 | 2.01 | <-4.38 | 0.0580 | 0.0253 | 0.106 | 0.155 | 0.314 | 0.302 | 0.486 | 0.0269 | 68.6 | 5.68 | <-8.09 | 0.418 | <-61.1 | 2.25 | 0.806 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV_50 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV_50 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 10:59:24 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | LFS_1/100 | 2.84 | 2.82 | 3.17 | 6.49 | 6.62 | >6.60 | 10.4 | 10.7 | 2.04 | 2.12 | 2.20 | 1.34 | 1.07 | 1.05 | 1.08 | 1.12 | 2343561 | 444802 | 271693 | 723946 | 471951 | >2021 | 1937 | 822 | 814 | 300 | 299 | 202 | 204 | 196 | 208 | 204 | 207 | 205 | 206 | 204 | 410 | 405 | 298 | 310 | 304 | >3045 | 3017 | 204 | 204 | 207 | 206 | 515 | 524 | 505 | 1022 | 947 | 2039 | 1993 | 198 | 150 | 201 | 206 | LFS_1/100 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | LFS_1/100 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | LFS_1/100 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | LFS_1/100 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | LFS_1/100 | 1
Mean | 02/07/2026 11:03:51 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Sis_Al+Fe | 0.138 | 0.134 | 0.279 | 3.28 | 0.0340 | 0.0276 | <0.735 | <1.14 | 0.006103 | 0.002344 | 0.132 | <-0.902 | 0.0639 | 0.0651 | 0.0594 | 0.121 | 2330770 | 353450 | 283639 | 607038 | 396452 | >8520 | 9283 | <-0.0130 | 9.53 | 17.3 | 3.37 | 1.22 | 1.08 | <-8.88 | <-0.0171 | 0.0417 | 0.212 | 0.649 | 0.766 | 3.90 | 0.599 | 1.46 | <-0.713 | 0.630 | 0.398 | >9690 | 9641 | 0.002689 | <-0.0320 | <-0.568 | 12.8 | 0.122 | 0.157 | <-12.0 | 3.21 | 87.0 | 5.50 | <-4.10 | 0.160 | <-61.3 | 8.68 | 6.70 | Sis_Al+Fe | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Sis_Al+Fe | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Sis_Al+Fe | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Sis_Al+Fe | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Sis_Al+Fe | 1
Mean | 02/07/2026 11:08:16 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_459_1/10 | <-529 | <-37.0 | 262 | 262 | 2.35 | 2.24 | <2.53 | <3.09 | 46.4 | 44.2 | 41.7 | 42.5 | >2.71 | 5.30 | 28.3 | 28.1 | 2334452 | 12040598 | 275259 | 634575 | 416383 | 5.92 | <-1.90 | 12.8 | 22.1 | 441 | 444 | 3.92 | 3.84 | 0.003503 | 0.0163 | 0.0115 | <-0.0213 | <-0.0652 | 1.12 | 0.960 | 0.203 | <-0.136 | 0.117 | 0.347 | 0.455 | 11.2 | 4.76 | 6.93 | 6.90 | 39.1 | 37.9 | 15.2 | 14.6 | 14.7 | <-0.136 | 65.3 | 33.7 | 11.7 | >2563 | 2736 | 0.953 | <-0.670 | AC_459_1/10 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | AC_459_1/10 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_459_1/10 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | AC_459_1/10 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_459_1/10 | 1
Mean | 02/07/2026 11:12:43 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | MA_54 | <-438 | <-47.9 | 143 | 145 | <-13.5 | <-18.2 | 25.5 | 26.1 | 25.3 | 24.1 | 23.0 | 23.8 | >2.70 | 5.35 | 64.6 | 63.4 | 2329033 | 25893635 | 271305 | 547395 | 361012 | <-0.0647 | <-7.67 | 6.54 | 14.4 | 235 | 234 | 45.9 | 46.9 | 0.009623 | 0.0279 | 0.005285 | <-0.0553 | <-0.005205 | 0.935 | 0.766 | 0.213 | 0.0306 | 1.11 | 0.990 | 1.12 | 5.64 | <-0.676 | 10.0 | 9.98 | 0.369 | 0.465 | 5.13 | 4.42 | 4.27 | 0.992 | 62.6 | 19.5 | <-9.18 | 524 | 501 | 19.2 | 18.2 | MA_54 | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | MA_54 | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | MA_54 | 1
Mean | 02/07/2026 11:17:09 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | MA_54_Dup | <-416 | <-42.7 | 141 | 140 | <-12.0 | <-16.2 | 25.1 | 25.5 | 25.2 | 23.7 | 22.5 | 23.4 | >2.61 | 5.18 | 63.8 | 62.6 | 2316664 | 25439252 | 278273 | 622803 | 407828 | <-0.389 | <-7.91 | 6.02 | 13.5 | 228 | 227 | 45.1 | 46.0 | 0.0184 | 0.002969 | 0.0161 | <-0.0285 | <-0.0221 | 0.973 | 0.781 | 0.215 | <-0.187 | 0.770 | 1.09 | 1.28 | 6.26 | <-0.0574 | 9.77 | 9.75 | 0.283 | 0.410 | 5.02 | 4.16 | 3.81 | 1.46 | 65.0 | 14.6 | <-14.7 | 510 | 489 | 20.5 | 19.3 | MA_54_Dup | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_Dup | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | MA_54_Dup | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_Dup | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | MA_54_Dup | 1
Mean | 02/07/2026 11:21:36 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | MA_54_ms_Na+k | <-618 | <-49.2 | 262 | 261 | <-282 | <-401 | 71.3 | 71.4 | 22.4 | 20.3 | 19.3 | 19.6 | >2.89 | 5.67 | 54.2 | 53.4 | 2323148 | 16684339 | 258982 | 541048 | 356990 | 43.1 | 34.0 | 5.03 | 14.8 | 212 | 213 | 39.3 | 39.0 | <-0.007391 | 0.0771 | <-0.0505 | 0.0739 | 0.0587 | 1.02 | 0.964 | 0.464 | 0.245 | 1.61 | 1.09 | 0.367 | 15.6 | 9.12 | 8.72 | 8.68 | 0.476 | 0.645 | 4.60 | 3.70 | 4.02 | 7.07 | 70.4 | 13.3 | <-11.5 | 432 | 397 | 24.7 | 23.7 | MA_54_ms_Na+k | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_ms_Na+k | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | MA_54_ms_Na+k | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_ms_Na+k | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | MA_54_ms_Na+k | 1
Mean | 02/07/2026 11:26:03 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | MA_54_ms_Mg+Ca | <-489 | <-74.9 | 140 | 142 | <-13.8 | <-17.1 | 25.7 | 26.1 | >81.6 | >80.7 | 75.8 | 79.0 | >2.94 | 5.91 | 167 | 167 | 2293998 | 17116213 | 241153 | 547885 | 359346 | 47.0 | 37.7 | 4.25 | 15.1 | 247 | 247 | 44.3 | 45.2 | <-0.003673 | 0.118 | <-0.0593 | 0.150 | 0.109 | 1.08 | 1.07 | 0.453 | 0.919 | 2.70 | 1.24 | 0.884 | 20.8 | 14.3 | 9.84 | 9.80 | 0.706 | 0.793 | 5.10 | 4.17 | 4.14 | 8.44 | 73.9 | 13.0 | <-15.7 | 499 | 475 | 37.0 | 35.1 | MA_54_ms_Mg+Ca | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_ms_Mg+Ca | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | MA_54_ms_Mg+Ca | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_ms_Mg+Ca | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | MA_54_ms_Mg+Ca | 1
Mean | 02/07/2026 11:30:30 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | MA_54_ms_B_300 ppb | <-421 | <-46.0 | 138 | 140 | <-12.3 | <-15.6 | 24.8 | 25.2 | 25.4 | 24.0 | 22.6 | 23.5 | >2.67 | 5.29 | 62.6 | 61.5 | 2333165 | 20793611 | 273267 | 646990 | 422835 | 283 | 267 | 300 | 305 | 525 | 531 | 320 | 320 | 280 | 297 | 291 | 281 | 278 | 276 | 276 | 281 | 277 | 280 | 279 | 285 | 286 | 277 | 374 | 374 | 281 | 272 | 283 | 287 | 284 | 271 | 352 | 319 | 289 | 768 | 769 | 300 | 305 | MA_54_ms_B_300 ppb | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_ms_B_300 ppb | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | MA_54_ms_B_300 ppb | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | MA_54_ms_B_300 ppb | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | MA_54_ms_B_300 ppb | 1
Mean | 02/07/2026 11:34:58 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Sis_Mg | 0.0917 | 0.0873 | 1.53 | 4.74 | 0.0702 | 0.0570 | <1.58 | <2.03 | >147 | >1192 | >1073 | >752 | 0.0767 | 0.0782 | 0.0819 | 0.172 | 2254767 | 947830 | 241727 | 588081 | 383320 | 4.09 | <-4.53 | 1.11 | 14.3 | <-0.466 | <-0.543 | 0.154 | 0.143 | <-0.005257 | 0.152 | <-0.0274 | 0.115 | 0.0139 | 1.10 | 1.43 | 0.0825 | 0.410 | 1.98 | 0.145 | 0.0754 | 2.71 | <-3.54 | 0.0289 | <-0.007628 | 0.137 | 0.132 | 0.308 | 0.233 | <-0.294 | <-0.679 | 68.6 | 6.20 | <-13.9 | 0.292 | <-61.2 | 4.48 | 2.76 | Sis_Mg | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Sis_Mg | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Sis_Mg | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Sis_Mg | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Sis_Mg | 1
Mean | 02/07/2026 11:39:25 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Sis_Ca | 0.0319 | 0.0274 | 1.17 | 3.51 | 0.0582 | 0.0577 | <1.44 | <1.96 | 0.146 | 0.132 | 0.267 | <-0.768 | >2.23 | 4.48 | >974 | >960 | 2291354 | 722668 | 261471 | 499532 | 331254 | 8.85 | <-17.4 | 1.77 | 14.7 | <-2.34 | <-1.85 | 0.365 | 0.385 | 0.0193 | 0.0382 | 0.0384 | <-0.0160 | 0.0469 | 0.727 | 0.762 | 0.688 | 0.674 | 1.85 | 1.41 | <-2.31 | 5.12 | <-1.42 | 0.413 | 0.361 | 1.09 | 1.09 | 0.413 | 0.474 | 0.964 | <-0.167 | 65.9 | 8.36 | <-4.51 | 6.46 | <-54.5 | 4.03 | 2.37 | Sis_Ca | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Sis_Ca | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | Sis_Ca | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Sis_Ca | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | Sis_Ca | 1
Mean | 02/07/2026 11:43:53 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | blank | <-0.0105 | <-0.0130 | 0.887 | 1.94 | 0.002037 | <-0.007189 | <1.14 | <1.59 | <-0.000158 | <-0.001053 | 0.130 | <-0.908 | 0.000966 | 0.000368 | <-0.001298 | 0.0769 | 2344336 | 505377 | 260815 | 722260 | 469814 | 0.119 | <-7.74 | 0.458 | 16.1 | 1.43 | 0.516 | <-0.0352 | <-0.143 | 0.0125 | 0.0296 | 0.0154 | 0.0445 | 0.000611 | 0.917 | 0.921 | 0.101 | 0.0789 | 0.463 | 0.0336 | 0.211 | <-0.101 | <-6.49 | <-0.006150 | <-0.0377 | <-0.0302 | 0.004046 | <-0.0500 | <-0.0313 | 0.104 | <-0.438 | 64.3 | 6.06 | <-5.97 | 0.005007 | <-61.5 | <-0.0177 | <-1.61 | blank | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | blank | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | blank | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | blank | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | blank | 1
Mean | 02/07/2026 11:48:18 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | blank | <-0.0108 | <-0.0141 | 0.710 | 1.92 | 0.001809 | <-0.007403 | <1.06 | <1.38 | <-0.000225 | <-0.001666 | 0.125 | <-0.908 | <-0.001725 | <-0.002465 | <-0.005544 | 0.0767 | 2349487 | 476991 | 246970 | 686053 | 447414 | 0.110 | <-7.81 | 0.126 | 15.2 | <-0.752 | <-1.14 | <-0.0147 | <-0.127 | <-0.004379 | 0.0802 | <-0.0278 | 0.0808 | <-0.0396 | 0.826 | 1.02 | 0.0956 | 0.272 | 1.08 | 0.0805 | 0.147 | <-0.147 | <-6.47 | 0.0234 | <-0.0126 | <-0.0112 | <-0.005029 | 0.002420 | <-0.131 | 0.101 | <-0.779 | 64.1 | 5.54 | <-5.80 | 0.0112 | <-61.5 | 0.0341 | <-1.60 | blank | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | blank | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | blank | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | blank | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | blank | 1
Mean | 02/07/2026 11:52:46 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | Blank | <-0.0116 | <-0.0142 | 0.563 | 2.40 | 0.001130 | <-0.008634 | <0.953 | <1.26 | <-0.000255 | <-0.001391 | 0.125 | <-0.908 | 0.000972 | 0.000351 | <-0.001270 | 0.0802 | 2341201 | 459296 | 243152 | 673388 | 437902 | 0.120 | <-8.37 | 1.40 | 14.0 | <-2.25 | <-2.15 | <-0.0139 | <-0.0913 | <-0.002762 | 0.0690 | <-0.0232 | 0.0459 | <-0.0674 | 0.800 | 0.961 | <-0.0590 | 0.309 | 1.46 | <-0.117 | 0.0369 | <-0.198 | <-6.42 | <-0.002142 | <-0.0378 | <-0.0199 | 0.006575 | 0.0749 | 0.266 | 0.262 | <-0.915 | 63.1 | 5.94 | <-6.42 | 0.006106 | <-61.5 | 0.0340 | <-1.59 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Blank | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | Blank | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 11:57:12 | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV 1 | 1.03 | 1.01 | 1.65 | 6.22 | 1.05 | 1.04 | <2.11 | <2.62 | 1.05 | 1.10 | 1.19 | 0.266 | 1.02 | 1.00 | 1.07 | 1.13 | 2354243 | 398609 | 268711 | 728902 | 476711 | 1083 | 1021 | 1009 | 1016 | 1000 | 1025 | 1031 | 1059 | 1015 | 1083 | 1062 | 1062 | 1053 | 1044 | 1031 | 1059 | 1041 | 1035 | 1040 | 1059 | 1041 | 1026 | 1014 | 1014 | 1068 | 1026 | 1050 | 1081 | 1050 | 1052 | 1029 | 1054 | 1017 | 1027 | 1053 | 1015 | 1044 |  | 1
MaxCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:33 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 12:01:39 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV_50 | <-40.8 | 1.16 | 46.8 | 51.2 | <-80.3 | <-112 | 49.6 | 49.6 | 53.6 | 54.1 | 50.0 | 52.4 | >2.83 | 5.55 | 50.8 | 50.6 | 2314759 | 402139 | 262549 | 717351 | 460530 | 14.3 | 5.04 | 1.76 | 14.2 | 5.57 | 5.64 | 0.413 | 0.365 | 0.0265 | 0.0850 | 0.0251 | 0.0338 | 0.0103 | 0.526 | 0.650 | 0.0740 | 0.004014 | 0.970 | 0.617 | 0.491 | 1.99 | <-4.36 | 0.0640 | 0.0305 | 0.0934 | 0.151 | 0.300 | 0.337 | 0.436 | 0.104 | 63.9 | 6.45 | <-3.30 | 0.397 | <-61.0 | 2.33 | 0.800 |  | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | CCV_50 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | CCV_50 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Mean | 02/07/2026 12:06:07 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | LFS_1/100 | 2.85 | 2.83 | 3.60 | 8.10 | 6.71 | >6.67 | 11.0 | 11.4 | 2.09 | 2.14 | 2.25 | 1.38 | 0.986 | 0.966 | 1.04 | 1.09 | 2340967 | 485342 | 267083 | 735801 | 474381 | >2075 | 1999 | 829 | 827 | 304 | 304 | 204 | 209 | 198 | 216 | 210 | 210 | 208 | 208 | 206 | 419 | 411 | 309 | 312 | 316 | >3095 | 3056 | 211 | 211 | 210 | 206 | 521 | 537 | 513 | 1034 | 950 | 2074 | 2026 | 203 | 158 | 204 | 207 | LFS_1/100 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | LFS_1/100 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | LFS_1/100 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | LFS_1/100 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | LFS_1/100 | 1
Mean | 02/07/2026 12:10:35 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | IHS_53_1/5 | <-560 | >7.22 | >1898 | >2078 | <-793 | <-1153 | 98.2 | 106 | >141 | >250 | 231 | 217 | >2.82 | 5.69 | 141 | 145 | 2276865 | 2846594 | 251723 | 501258 | 328978 | 0.510 | <-9.23 | 2.66 | 14.7 | 935 | 961 | 8.91 | 8.49 | 0.0140 | 0.0779 | <-0.0263 | <-0.0767 | 0.0697 | 0.533 | 0.270 | 0.215 | 0.175 | 2.26 | 1.94 | 1.49 | 0.627 | <-5.65 | 57.3 | 57.3 | 15.8 | 16.3 | 3.12 | 2.86 | 3.60 | 0.546 | 62.2 | 10.1 | 1.46 | >1744 | 1838 | 102 | 96.3 | IHS_53_1/5 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_53_1/5 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | IHS_53_1/5 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_53_1/5 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | IHS_53_1/5 | 1
Mean | 02/07/2026 12:15:04 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | IHS_54_1/2 | <-699 | <-14.2 | >1324 | >1364 | <-163 | <-192 | 55.1 | 61.8 | >157 | >188 | 174 | 171 | >3.02 | >6.09 | 173 | 178 | 2276137 | 18407266 | 234974 | 529849 | 342369 | <-0.306 | <-8.38 | 5.53 | 12.3 | 704 | 724 | 52.0 | 48.9 | 0.002272 | 0.137 | <-0.0732 | <-0.0406 | 0.0357 | 0.922 | 0.786 | 0.285 | 0.493 | 4.22 | 2.87 | 1.56 | 0.516 | <-5.84 | 34.7 | 34.7 | 259 | 249 | 3.83 | 2.04 | 2.12 | 0.894 | 68.2 | 15.0 | <-13.5 | >4295 | 4819 | 65.0 | 62.0 | IHS_54_1/2 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_54_1/2 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | IHS_54_1/2 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_54_1/2 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | IHS_54_1/2 | 1
Mean | 02/07/2026 12:19:32 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | IHS_55_1/5 | <-670 | <-15.8 | >1177 | >1203 | <-74.6 | <-88.4 | 39.0 | 43.2 | >152 | >169 | 153 | 154 | >3.00 | >6.01 | 128 | 130 | 2293163 | 6617480 | 242761 | 465372 | 307715 | 0.0281 | <-8.86 | 2.55 | 15.1 | 378 | 387 | 25.9 | 25.6 | 0.0108 | 0.120 | <-0.0818 | <-0.0394 | 0.0139 | 0.601 | 0.469 | 0.0976 | 0.376 | 4.61 | 3.25 | 2.77 | 0.0704 | <-6.33 | 22.3 | 22.3 | 244 | 234 | 1.95 | 1.25 | 1.53 | 1.01 | 64.1 | 13.6 | <-4.98 | >3382 | 3684 | 62.4 | 59.6 | IHS_55_1/5 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_55_1/5 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | IHS_55_1/5 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | IHS_55_1/5 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | IHS_55_1/5 | 1
Mean | 02/07/2026 12:23:58 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_455_1/5 | <-592 | <-14.8 | >1004 | >1008 | 6.28 | >6.06 | 11.6 | 14.1 | >139 | >144 | 134 | 134 | >2.95 | 5.77 | 50.8 | 51.4 | 2293376 | 23583712 | 254138 | 610177 | 397303 | 1.73 | <-6.57 | 11.1 | 18.2 | 2723 | 2809 | 5.34 | 5.15 | <-0.005540 | 0.101 | <-0.0432 | <-0.0923 | 0.0421 | 0.367 | 0.274 | 0.573 | 0.735 | 1.09 | 0.304 | 0.155 | 15.3 | 8.59 | 26.5 | 26.5 | 17.9 | 17.7 | 7.61 | 5.95 | 6.24 | 1.12 | 66.9 | 15.2 | <-15.0 | >1737 | 1834 | 1.01 | <-0.668 | AC_455_1/5 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_455_1/5 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_455_1/5 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_455_1/5 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_455_1/5 | 1
Mean | 02/07/2026 12:28:24 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_456_1/5 | <-674 | <-30.5 | >789 | >797 | 6.20 | >6.02 | 8.83 | 9.91 | >133 | >130 | 121 | 121 | >2.96 | 5.95 | 129 | 131 | 2288387 | 13513013 | 242336 | 617619 | 398791 | 3.25 | <-9.20 | 62.5 | 69.5 | 2019 | 2072 | 4.01 | 3.94 | <-2.99 | 0.139 | <-0.0604 | 0.0488 | 0.241 | 0.283 | 1.13 | 0.240 | 0.976 | 1.09 | 0.392 | <-0.953 | >3538 | 3481 | 26.7 | 26.7 | 1874 | 1801 | 3.26 | 2.48 | <-3.32 | 0.746 | 77.7 | 12.4 | <-12.3 | >2436 | 2610 | 31.6 | 29.1 | AC_456_1/5 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_456_1/5 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_456_1/5 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_456_1/5 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_456_1/5 | 1
Mean | 02/07/2026 12:32:50 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_457_1/5 | <-564 | <-9.98 | >1088 | >1104 | 6.36 | >6.13 | 9.35 | 11.0 | >139 | >153 | 140 | 139 | >2.73 | 5.50 | 132 | 134 | 2278161 | 6956517 | 259575 | 493988 | 325266 | 3.97 | <-5.44 | 4.07 | 15.7 | 2070 | 2134 | 4.39 | 4.25 | 0.0101 | 0.0698 | <-0.0241 | <-0.0710 | <-0.0136 | 0.525 | 0.426 | 0.163 | <-0.0779 | 0.897 | 0.426 | <-0.800 | 1.75 | <-4.69 | 36.8 | 36.7 | 602 | 574 | 9.61 | 9.48 | 10.2 | 0.909 | 67.2 | 10.7 | <-1.21 | >2173 | 2318 | 18.6 | 16.0 | AC_457_1/5 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_457_1/5 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_457_1/5 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_457_1/5 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_457_1/5 | 1
Mean | 02/07/2026 12:37:17 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_458_1/5 | <-584 | <-4.65 | >1368 | >1430 | 7.06 | >6.90 | 10.9 | 12.5 | >142 | >173 | 159 | 158 | >2.77 | 5.59 | 150 | 154 | 2264059 | 14404404 | 257761 | 284357 | 195124 | 0.278 | <-8.98 | 6.31 | 15.3 | 1915 | 1971 | 6.33 | 6.25 | 0.0133 | 0.0874 | <-0.0327 | <-0.113 | 0.0494 | 2.72 | 2.46 | 0.379 | 0.151 | 0.981 | 0.921 | <-0.544 | 3.13 | <-3.55 | 46.1 | 46.1 | 1473 | 1403 | 39.9 | 39.7 | 40.9 | 0.588 | 63.4 | 16.3 | <-7.77 | >2616 | 2818 | 4.12 | 2.19 | AC_458_1/5 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_458_1/5 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_458_1/5 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_458_1/5 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_458_1/5 | 1
Mean | 02/07/2026 12:41:43 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_459_1/2 | <-652 | <-8.43 | >1398 | >1471 | 6.00 | 5.98 | 13.2 | 14.4 | >149 | >242 | 230 | 218 | >2.95 | 5.93 | 154 | 161 | 2213126 | 53588440 | 242206 | 538073 | 344383 | 29.4 | 23.2 | 48.5 | 46.3 | 2406 | 2475 | 20.7 | 20.1 | <-0.0221 | 0.178 | <-0.103 | <-0.0948 | <-0.0203 | 4.54 | 4.18 | 0.835 | 0.735 | 1.71 | 1.90 | 0.389 | 69.9 | 62.3 | 38.4 | 38.4 | 211 | 204 | 80.5 | 75.7 | 78.2 | 0.318 | 67.9 | 54.7 | <-7.90 | >4504 | >12994 | 1.96 | <-0.209 | AC_459_1/2 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_459_1/2 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_459_1/2 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_459_1/2 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_459_1/2 | 1
Mean | 02/07/2026 12:46:11 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | AC_460_1/2 | <-691 | <-36.0 | >714 | >709 | <-2695 | <-4959 | >249 | >251 | >151 | >194 | 177 | 180 | >2.79 | 5.62 | >312 | >320 | 2236307 | 24025143 | 243972 | 447297 | 292501 | 4.78 | <-7.50 | 8.18 | 16.3 | 1835 | 1887 | 12.1 | 12.4 | 0.004776 | 0.161 | <-0.0939 | <-0.0508 | <-0.0294 | 0.857 | 0.969 | 0.181 | 0.302 | 1.38 | 0.367 | <-1.06 | 6.38 | <-0.322 | 82.8 | 82.8 | 1843 | 1765 | 76.8 | 72.6 | 75.2 | 0.305 | 65.0 | 32.0 | <-4.94 | >4494 | 5871 | 17.3 | 14.6 | AC_460_1/2 | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_460_1/2 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | AC_460_1/2 | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | AC_460_1/2 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | AC_460_1/2 | 1
Mean | 02/07/2026 12:50:38 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.0472 | 0.0366 | 16.0 | 4.77 | 0.0147 | 0.004962 | <4.76 | 6.11 | 0.003381 | 0.002451 | 0.122 | <-0.905 | 0.007394 | 0.006492 | 0.007449 | 0.0909 | 2322167 | 1499437 | 249239 | 673769 | 434096 | 0.0763 | <-8.45 | 0.363 | 15.2 | 6.66 | 6.39 | <-0.0402 | <-0.115 | 0.0102 | 0.0858 | <-0.0163 | 0.009640 | <-0.009617 | 0.503 | 0.492 | <-0.0824 | 0.229 | 0.671 | <-0.184 | 0.117 | <-0.182 | <-6.52 | <-0.0195 | <-0.0546 | 0.0173 | 0.0508 | <-0.0525 | 0.0614 | 0.156 | 0.315 | 63.5 | 10.8 | <-1.29 | 0.175 | <-61.3 | 0.0213 | <-1.59 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 12:55:04 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.0195 | 0.0137 | 11.7 | 3.43 | 0.007158 | <-0.001652 | <3.37 | <4.19 | 0.001537 | 0.000144 | 0.125 | <-0.906 | 0.001775 | 0.001189 | <-0.000566 | 0.0829 | 2348412 | 949081 | 246573 | 673168 | 437723 | 0.0941 | <-7.89 | 0.598 | 14.6 | 1.96 | 1.07 | 0.007088 | <-0.163 | 0.008615 | 0.0808 | <-0.0231 | 0.0232 | <-0.0762 | 0.382 | 0.616 | 0.0188 | 0.166 | 1.09 | 0.006859 | 0.0531 | <-0.148 | <-6.41 | 0.008617 | <-0.0284 | <-0.003131 | 0.0280 | 0.0268 | 0.170 | 0.0862 | <-0.141 | 65.2 | 8.81 | <-5.09 | 0.0592 | <-61.4 | 0.0181 | <-1.57 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 12:59:30 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.133 | 0.124 | 54.8 | 21.7 | 0.0958 | 0.009551 | 14.7 | 11.6 | 0.0120 | 0.005801 | 0.0983 | <-0.901 | 0.0505 | 0.0513 | 0.151 | 0.245 | 2358006 | 839725 | 114314 | 670217 | 438610 | 1.34 | <-2.79 | <-14.2 | 21.0 | 8.60 | 8.78 | <-1.19 | <-0.307 | <-0.266 | 3.37 | <-1.99 | 2.25 | <-0.297 | <-1.38 | 1.67 | 0.462 | 13.4 | 36.6 | <-4.87 | <-4.62 | <-1.15 | <-6.46 | 0.218 | 0.182 | 0.554 | 0.103 | 1.57 | 1.69 | <-4.92 | <-6.62 | 104 | <-19.3 | <-172 | 0.450 | <-61.3 | 2.10 | <-1.85 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 13:03:56 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.005206 | 0.000104 | 7.45 | 2.78 | 0.001326 | <-0.005834 | <2.27 | <2.92 | 0.000054 | <-0.001192 | 0.115 | <-0.908 | <-0.004384 | <-0.005199 | <-0.008222 | 0.0802 | 2335862 | 554075 | 274921 | 753343 | 485422 | 0.141 | <-8.15 | 1.25 | 13.3 | <-0.705 | <-0.979 | <-0.0129 | <-0.116 | 0.003486 | <-0.0167 | 0.0176 | <-0.0104 | 0.004064 | 0.466 | 0.527 | 0.0168 | <-0.266 | <-0.147 | 0.126 | 0.429 | 0.004100 | <-6.34 | <-0.0325 | <-0.0707 | <-0.006634 | 0.0145 | <-0.0736 | 0.0603 | 0.543 | 0.0372 | 64.4 | 8.50 | <-2.18 | 0.003557 | <-61.5 | <-0.113 | <-1.63 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 13:08:23 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.002739 | <-0.001917 | 6.68 | 6.36 | 0.001304 | <-0.006742 | <2.07 | <2.69 | 0.000078 | <-0.000482 | 0.139 | <-0.908 | <-0.004336 | <-0.005152 | <-0.005014 | 0.0767 | 2352162 | 492153 | 269979 | 730719 | 473578 | 0.190 | <-8.11 | <-0.0526 | 14.4 | <-1.45 | <-1.35 | <-0.0118 | <-0.0861 | 0.008137 | 0.008510 | 0.0461 | 0.0182 | <-0.0379 | 0.554 | 0.420 | 0.0676 | <-0.226 | 0.317 | 0.204 | 0.374 | <-0.0520 | <-6.36 | <-0.009366 | <-0.0431 | <-0.0153 | 0.0179 | <-0.0658 | <-0.0878 | <-0.0473 | <-0.231 | 64.3 | 8.05 | <-6.72 | 0.005369 | <-61.4 | <-0.133 | <-1.60 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 13:12:49 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.001389 | <-0.002053 | 5.95 | 1.90 | 0.000482 | <-0.008886 | <1.91 | <2.46 | 0.000053 | <-0.000246 | 0.129 | <-0.907 | <-0.004252 | <-0.005076 | <-0.004847 | 0.0765 | 2345148 | 446996 | 265339 | 728271 | 470745 | 0.200 | <-8.47 | 1.18 | 13.4 | <-0.950 | <-1.43 | 0.0342 | <-0.0771 | 0.006692 | 0.0402 | <-0.002072 | 0.0109 | <-0.0634 | 0.308 | 0.501 | 0.200 | <-0.153 | 0.154 | 0.0636 | 0.631 | 0.317 | <-6.11 | <-0.0138 | <-0.0490 | 0.002420 | 0.0235 | 0.0516 | 0.168 | 0.395 | <-0.188 | 63.5 | 7.52 | <-1.45 | 0.001003 | <-61.4 | <-0.121 | <-1.57 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 13:17:15 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | 0.000407 | <-0.003833 | 5.39 | 3.64 | <-0.000026 | <-0.007493 | <1.78 | <2.31 | 0.000128 | <-0.001156 | 0.127 | <-0.908 | <-0.004113 | <-0.004916 | <-0.005721 | 0.0828 | 2349925 | 419146 | 260093 | 722255 | 466589 | 0.242 | <-8.86 | 0.503 | 13.8 | <-0.825 | <-2.05 | 0.177 | 0.0148 | <-0.000432 | 0.0349 | 0.0173 | 0.0179 | 0.0102 | 0.337 | 0.486 | <-0.0278 | 0.000854 | 0.569 | <-0.0856 | 0.611 | 0.269 | <-6.02 | <-0.0133 | <-0.0496 | 0.009749 | 0.002073 | <-0.0316 | 0.004248 | 0.259 | <-0.009142 | 64.5 | 5.58 | <-4.69 | 0.008158 | <-61.4 | <-0.0539 | <-1.63 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 13:21:40 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | BLANK | <-0.000611 | <-0.004933 | 4.86 | 1.64 | 0.000808 | <-0.006934 | <1.67 | <2.06 | 0.000144 | <-0.000607 | 0.159 | <-0.908 | <-0.003992 | <-0.004779 | <-0.004009 | 0.0862 | 2360193 | 407038 | 255644 | 709268 | 460974 | 0.221 | <-7.28 | 1.11 | 14.1 | 5.85 | 5.84 | 0.651 | 0.419 | 0.006347 | 0.0389 | <-0.0140 | 0.0678 | <-0.003709 | 0.368 | 0.383 | <-0.0202 | 0.0850 | 0.658 | 0.326 | 0.0495 | 0.362 | <-5.93 | <-0.0117 | <-0.0483 | 0.007410 | 0.0334 | <-0.106 | <-0.0561 | 0.0746 | <-0.145 | 64.3 | 6.14 | <-2.52 | 0.005427 | <-61.5 | <-0.0173 | <-1.56 | BLANK | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 | BLANK | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | BLANK | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | BLANK | 1
Mean | 02/07/2026 13:26:07 | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  | 1 | Conc | CCV 1 | 1.01 | 0.986 | 4.25 | 6.12 | 1.05 | 1.04 | <2.31 | <2.79 | 1.05 | 1.10 | 1.17 | 0.248 | 1.01 | 1.00 | 1.08 | 1.14 | 2351158 | 480500 | 271349 | 737653 | 480425 | 1096 | 1019 | 994 | 1018 | 999 | 1030 | 1035 | 1043 | 1010 | 1083 | 1057 | 1063 | 1064 | 1043 | 1033 | 1069 | 1048 | 1027 | 1036 | 1064 | 1048 | 1023 | 1014 | 1014 | 1068 | 1019 | 1045 | 1089 | 1048 | 1044 | 1018 | 1051 | 1021 | 1012 | 1056 | 1017 | 1050 |  | 1
MaxCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 5.00 | 5.00 | 500 | 500 | 10.00 | 6.00 | 240 | 200 | 60.0 | 60.0 | 240 | 240 | 1.20 | 6.00 | 240 | 240 | -- | -- | -- | -- | -- | 1200 | 12000 | 6000 | 12000 | 12000 | 12000 | 12000 | 12000 | 12000 | 1200 | 12000 | 1200 | 12000 | 12000 | 6000 | 6000 | 12000 | 12000 | 1200 | 1200 | 1200 | 12000 | 12000 | 12000 | 12000 | 6000 | 1200 | 12000 | 12000 | 1200 | 12000 | 6000 | 12000 | 1200 | 12000 | 6000 | 12000 |  | 1
MinCalib |  | 06/07/2026 13:38:34 | Mjr_Metals_20241201 |  |  |  | Conc | CCV 1 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 5.00 | 5.00 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | -- | -- | -- | -- | -- | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 | 0.000000 |  | 1
Note on J1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: כל עמודה מייצגת פרמטר אחד Reply: J K L זה פרמטר אחד שיש לו כמה ערכים. צריך לעשות עבודה לפי כללים מסויימים מה מהם נכנס (לפי איזו עקומה להתייחס)

## Method Properties
Name | Mjr_Metals_20241201
Description
Created by
Creation Date | 15/03/2023 13:42:15
Last Modification | 06/07/2026 13:38:33
Version | 20.1.19.1.47.1.61.1.17.1.2.1.5.1.1
Design For | FMT46
Plasma View | EOP / SOP
Method Type | Normal
Solvent Type | Water
Nebulizer Type | Crossflow

## Method Measurement Conditions
Condition Name | SOP (1) | Wavelength Range [nm] | 285.9 - 770.6
Sample Introduction [ml/min] | 2
Stabilization Time [s] | 0
Strategy Type | Standard
Interval Time [ms] | Measure Time [ms]
1 | 1000
10 | 1000
100 | 2000
1000 | 4000
View Type | SOP
Plasma Power [W] | 1300
Gas Flows [ml/min]
Auxiliary | 800
Coolant | 14000
Nebulizer | 800
Condition Name | EOP (2) | Wavelength Range [nm] | 163 - 285.7
Sample Introduction [ml/min] | 2
Stabilization Time [s] | 0
Strategy Type | Standard
Interval Time [ms] | Measure Time [ms]
1 | 1000
10 | 1000
100 | 2000
1000 | 4000
5000 | 20000
View Type | EOP
Plasma Power [W] | 1300
Gas Flows [ml/min]
Auxiliary | 1600
Coolant | 14000
Nebulizer | 800
Condition Name | EOP (1) | Wavelength Range [nm] | 285.9 - 770.6
Sample Introduction [ml/min] | 2
Stabilization Time [s] | 0
Strategy Type | Fast
Interval Time [ms] | Measure Time [ms]
1 | 500
10 | 500
100 | 1000
View Type | EOP
Plasma Power [W] | 1300
Gas Flows [ml/min]
Auxiliary | 1600
Coolant | 14000
Nebulizer | 800
Condition Name | SOP (2) | Wavelength Range [nm] | 163 - 285.7
Sample Introduction [ml/min] | 2
Stabilization Time [s] | 0
Strategy Type | Fast
Interval Time [ms] | Measure Time [ms]
1 | 500
10 | 500
100 | 1000
1000 | 1000
View Type | SOP
Plasma Power [W] | 1300
Gas Flows [ml/min]
Auxiliary | 800
Coolant | 14000
Nebulizer | 800

## Method Standard Definitions
Standard Name | As [ppb] | B [ppb] | Ba [ppb] | Be [ppb] | C [ppb] | Ca [ppm] | Cd [ppb] | Co [ppb] | Cr [ppb] | Cu [ppb] | Fe [ppb] | K [ppm] | Li [ppb] | Mg [ppm] | Mn [ppb] | Na [ppm] | Ni [ppb] | Pb [ppb] | Se [ppb] | Sr [ppb] | Zn [ppb] | Al [ppb]
0.01 | 10 | 10 | 10 | 10 | 10 | 0.01 | 10 | 10 | 10 | 10 | 10 | 0.01 | 10 | 0.01 | 10 | 0.01 | 10 | 10 | 10 | 10 | 10 | 10
0.05 | 50 | 50 | 50 | 50 | -- | 0.05 | 50 | 50 | 50 | 50 | 50 | 0.05 | 50 | 0.05 | 50 | 0.05 | 50 | 50 | 50 | 50 | 50 | 50
0.1 | 100 | 100 | 100 | 100 | -- | 0.1 | 100 | 100 | 100 | 100 | 100 | 0.1 | 100 | 0.1 | 100 | 0.1 | 100 | 100 | 100 | 100 | 100 | 100
0.5 | 500 | 500 | 500 | 500 | -- | 0.5 | 500 | 500 | 500 | 500 | 500 | 0.5 | 500 | 0.5 | 500 | 0.5 | 500 | 500 | 500 | 500 | 500 | 500
1 | 1000 | 1000 | 1000 | 1000 | -- | 1 | 1000 | 1000 | 1000 | 1000 | 1000 | 1 | 1000 | 1 | 1000 | 1 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
5 | 5000 | 5000 | 5000 | 5000 | -- | 5 | 5000 | 5000 | 5000 | 5000 | 5000 | 5 | 5000 | 5 | 5000 | 5 | 5000 | 5000 | 5000 | 5000 | 5000 | 5000
10 | 10000 | 10000 | 10000 | 10000 | -- | 10 | 10000 | 10000 | 10000 | 10000 | 10000 | 10 | 10000 | 10 | 10000 | 10 | 10000 | 10000 | 10000 | 10000 | 10000 | 10000
50 | -- | -- | -- | -- | -- | 50 | -- | -- | -- | -- | -- | 50 | -- | 50 | -- | 50 | -- | -- | -- | -- | -- | --
100 | -- | -- | -- | -- | -- | 100 | -- | -- | -- | -- | -- | 100 | -- | 100 | -- | 100 | -- | -- | -- | -- | -- | --
200 | -- | -- | -- | -- | -- | 200 | -- | -- | -- | -- | -- | 200 | -- | 200 | -- | 200 | -- | -- | -- | -- | -- | --
Blank | 0 | 0 | 0 | 0 | -- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Method Line Definitions
Line Name | Display Unit | Line Type | Region From [nm] | Region To [nm] | Peak Position [nm] | Offset [nm] | Background Correction Mode | Background Polynomial Order | Fix Background Peak Height | Left Background From [nm] | Left Background To [nm] | Right Background From [nm] | Right Background To [nm] | Integration Mode | Line Range From [nm] | Line Range To [nm] | Peak Fit Fix FWHM [nm] | Peak Fit Fix Peak Position [nm] | Line Switch | Reference Line
Na 589.592 | ppm | Analytical | 589.494 | 589.727 | 589.593 | 0.001 | Polynomial | 0 |  | 589.512 | 589.514 | 589.672 | 589.676 | Range Integration | 589.592 | 589.593 |  |  | True | Y 224.306
Na 588.995 | ppm | Analytical | 588.845 | 589.145 | 588.995 | 0 | Polynomial | 2 |  | 588.947 | 588.952 | 589.053 | 589.056 | Range Integration | 588.994 | 588.996 |  |  | True | Y 224.306
Na 589.592 R | ppm | Analytical | 589.443 | 589.743 | 589.593 | 0.001 | Polynomial | 0 |  | 589.522 | 589.524 | 589.676 | 589.682 | Range Integration | 589.592 | 589.594 |  |  | True | Y 224.306
Na 330.298 R | ppm | Analytical | 330.146 | 330.446 | 330.296 | -0.002 | Polynomial | 2 |  | 330.276 | 330.277 | 330.331 | 330.338 | Range Integration | 330.295 | 330.297 |  |  | True | Y 224.306
K 766.491 | ppm | Analytical | 766.344 | 766.644 | 766.494 | 0.003 | Polynomial | 1 |  | 766.451 | 766.452 | 766.524 | 766.526 | Range Integration | 766.493 | 766.495 |  |  | True | Y 224.306
K 769.896 | ppm | Analytical | 769.826 | 770.001 | 769.899 | 0.003 | Polynomial | 1 |  | 769.86 | 769.866 | 769.953 | 769.959 | Range Integration | 769.898 | 769.9 |  |  | True | Y 224.306
K 766.491 R | ppm | Analytical | 766.344 | 766.644 | 766.494 | 0.003 | Polynomial | 1 |  | 766.453 | 766.454 | 766.546 | 766.546 | Range Integration | 766.493 | 766.495 |  |  | True | Y 224.306
K 769.896 R | ppm | Analytical | 769.749 | 770.049 | 769.899 | 0.003 | Polynomial | 1 |  | 769.854 | 769.861 | 769.942 | 769.944 | Range Integration | 769.898 | 769.9 |  |  | True | Y 224.306
Mg 285.213 | ppm | Analytical | 285.069 | 285.369 | 285.219 | 0.006 | Polynomial | 0 |  | 285.17 | 285.172 | 285.317 | 285.323 | Range Integration | 285.218 | 285.22 |  |  | True | Y 224.306
Mg 279.079 | ppm | Analytical | 278.923 | 279.223 | 279.073 | -0.006 | Polynomial | 0 |  | 279.011 | 279.013 | 279.138 | 279.142 | Range Integration | 279.072 | 279.074 |  |  | True | Y 224.306
Mg 279.079 R | ppm | Analytical | 278.923 | 279.223 | 279.073 | -0.006 | Polynomial | 0 |  | 279.024 | 279.025 | 279.11 | 279.111 | Range Integration | 279.072 | 279.074 |  |  | True | Y 224.306
Mg 285.213 R | ppm | Analytical | 285.069 | 285.369 | 285.219 | 0.006 | Polynomial | 0 |  | 285.158 | 285.162 | 285.273 | 285.277 | Range Integration | 285.218 | 285.22 |  |  | True | Y 224.306
Ca 393.366 | ppm | Analytical | 393.217 | 393.517 | 393.367 | 0.001 | Polynomial | 0 |  |  |  | 393.456 | 393.459 | Range Integration | 393.366 | 393.368 |  |  | True | Y 224.306
Ca 396.847 | ppm | Analytical | 396.699 | 396.999 | 396.849 | 0.002 | Polynomial | 0 |  | 396.768 | 396.769 | 396.915 | 396.916 | Range Integration | 396.848 | 396.85 |  |  | True | Y 224.306
Ca 317.933 | ppm | Analytical | 317.784 | 318.084 | 317.934 | 0.001 | Polynomial | 0 |  | 317.854 | 317.857 | 318.019 | 318.02 | Range Integration | 317.933 | 317.935 |  |  | True | Y 224.306
Ca 422.673 | ppm | Analytical | 422.525 | 422.825 | 422.675 | 0.002 | Polynomial | 0 |  | 422.603 | 422.605 | 422.75 | 422.753 | Range Integration | 422.674 | 422.676 |  |  | True | Y 224.306
Ar 404.442 | - | Monitor | 404.294 | 404.594 | 404.444 | 0.002 | None | [6 empty] | Range Integration | 404.443 | 404.445
C 193.091 | - | Monitor | 192.941 | 193.241 | 193.091 | 0 | None | [6 empty] | Range Integration | 193.09 | 193.092
Y 224.306 | - | Reference | 224.156 | 224.456 | 224.306 | 0 | Polynomial | 0 |  | 224.292 | 224.292 | 224.318 | 224.319 | Range Integration | 224.305 | 224.307
Y 371.030 | - | Reference | 370.88 | 371.18 | 371.03 | 0 | None | [6 empty] | Range Integration | 371.029 | 371.031
Y 324.228 | - | Reference | 324.08 | 324.38 | 324.23 | 0.002 | None | [6 empty] | Range Integration | 324.229 | 324.231
Al 167.078 | ppb | Analytical | 166.927 | 167.227 | 167.077 | -0.001 | Polynomial | 0 |  | 167.026 | 167.027 | 167.133 | 167.134 | Range Integration | 167.076 | 167.078 |  |  | True | Y 224.306
Al 396.152 | ppb | Analytical | 396.005 | 396.305 | 396.155 | 0.003 | Polynomial | 1 |  | 396.069 | 396.072 | 396.245 | 396.249 | Range Integration | 396.154 | 396.156 |  |  | True | Y 224.306
As 189.042 | ppb | Analytical | 188.894 | 189.194 | 189.044 | 0.002 | Polynomial | 0 |  | 189.014 | 189.016 | 189.072 | 189.074 | Range Integration | 189.043 | 189.045 |  |  | True | Y 224.306
As 193.759 | ppb | Analytical | 193.611 | 193.911 | 193.761 | 0.002 | Polynomial | 0 |  | 193.734 | 193.736 | 193.786 | 193.787 | Range Integration | 193.76 | 193.762 |  |  | True | Y 224.306
B 249.773 | ppb | Analytical | 249.624 | 249.924 | 249.774 | 0.001 | Polynomial | 0 |  | 249.736 | 249.739 | 249.8 | 249.803 | Range Integration | 249.773 | 249.775 |  |  | True | Y 224.306
B 249.677 | ppb | Analytical | 249.528 | 249.828 | 249.678 | 0.001 | Polynomial | 0 |  | 249.658 | 249.659 | 249.696 | 249.696 | Range Integration | 249.677 | 249.679 |  |  | True | Y 224.306
Ba 233.527 | ppb | Analytical | 233.378 | 233.678 | 233.528 | 0.001 | Polynomial | 0 |  | 233.49 | 233.491 | 233.563 | 233.564 | Range Integration | 233.527 | 233.529 |  |  | True | Y 224.306
Ba 455.404 | ppb | Analytical | 455.258 | 455.558 | 455.408 | 0.004 | Polynomial | 0 |  | 455.366 | 455.367 | 455.452 | 455.452 | Range Integration | 455.407 | 455.409 |  |  | True | Y 224.306
Be 234.861 | ppb | Analytical | 234.711 | 235.011 | 234.861 | 0 | Polynomial | 1 |  | 234.838 | 234.84 | 234.902 | 234.904 | Range Integration | 234.86 | 234.862 |  |  | True | Y 224.306
Be 313.042 | ppb | Analytical | 312.894 | 313.194 | 313.044 | 0.002 | Polynomial | 0 |  | 312.975 | 312.976 | 313.079 | 313.081 | Range Integration | 313.043 | 313.045 |  |  | True | Y 224.306
Be 313.107 | ppb | Analytical | 312.959 | 313.259 | 313.109 | 0.002 | Polynomial | 1 |  | 313.077 | 313.079 | 313.157 | 313.16 | Range Integration | 313.108 | 313.11 |  |  | True | Y 224.306
Cd 214.438 | ppb | Analytical | 214.29 | 214.59 | 214.44 | 0.002 | Polynomial | 0 |  | 214.415 | 214.416 | 214.489 | 214.489 | Range Integration | 214.439 | 214.441 |  |  | True | Y 224.306
Cd 226.502 | ppb | Analytical | 226.351 | 226.651 | 226.501 | -0.001 | Polynomial | 0 |  | 226.468 | 226.469 | 226.528 | 226.529 | Range Integration | 226.5 | 226.502 |  |  | True | Y 224.306
Co 228.616 | ppb | Analytical | 228.465 | 228.765 | 228.615 | -0.001 | Polynomial | 0 |  | 228.588 | 228.588 | 228.656 | 228.657 | Range Integration | 228.614 | 228.616 |  |  | True | Y 224.306
Co 238.892 | ppb | Analytical | 238.741 | 239.041 | 238.891 | -0.001 | Polynomial | 0 |  | 238.803 | 238.805 | 238.928 | 238.93 | Range Integration | 238.89 | 238.892 |  |  | True | Y 224.306
Cr 267.716 | ppb | Analytical | 267.567 | 267.867 | 267.717 | 0.001 | Polynomial | 0 |  | 267.676 | 267.677 | 267.747 | 267.748 | Range Integration | 267.716 | 267.718 |  |  | True | Y 224.306
Cr 283.563 | ppb | Analytical | 283.416 | 283.716 | 283.566 | 0.003 | Polynomial | 0 |  | 283.531 | 283.533 | 283.588 | 283.588 | Range Integration | 283.565 | 283.567 |  |  | True | Y 224.306
Cu 324.754 | ppb | Analytical | 324.607 | 324.907 | 324.757 | 0.003 | Polynomial | 0 |  | 324.679 | 324.685 | 324.818 | 324.823 | Range Integration | 324.756 | 324.758 |  |  | True | Y 224.306
Cu 224.700 | ppb | Analytical | 224.553 | 224.853 | 224.703 | 0.003 | Polynomial | 0 |  | 224.68 | 224.681 | 224.717 | 224.719 | Range Integration | 224.702 | 224.704 |  |  | True | Y 224.306
Cu 327.396 | ppb | Analytical | 327.245 | 327.545 | 327.395 | -0.001 | Polynomial | 0 |  | 327.321 | 327.323 | 327.448 | 327.449 | Range Integration | 327.394 | 327.396 |  |  | True | Y 224.306
Fe 238.204 | ppb | Analytical | 238.054 | 238.354 | 238.204 | 0 | Polynomial | 0 |  | 238.186 | 238.187 | 238.225 | 238.227 | Range Integration | 238.203 | 238.205 |  |  | True | Y 224.306
Fe 259.941 | ppb | Analytical | 259.79 | 260.09 | 259.94 | -0.001 | Polynomial | 0 |  | 259.91 | 259.911 | 259.976 | 259.978 | Range Integration | 259.939 | 259.941 |  |  | True | Y 224.306
Li 670.780 | ppb | Analytical | 670.628 | 670.928 | 670.778 | -0.002 | Polynomial | 0 |  | 670.628 | 670.634 | 670.922 | 670.928 | Range Integration | 670.777 | 670.779 |  |  | True | Y 224.306
Li 670.780 | ppb | Analytical | 670.628 | 670.928 | 670.778 | -0.002 | Polynomial | 0 |  | 670.628 | 670.63 | 670.921 | 670.928 | Range Integration | 670.777 | 670.779 |  |  | True | Y 224.306
Mn 257.611 | ppb | Analytical | 257.461 | 257.761 | 257.611 | 0 | Polynomial | 0 |  | 257.574 | 257.576 | 257.656 | 257.658 | Range Integration | 257.61 | 257.612 |  |  | True | Y 224.306
Mn 259.373 | ppb | Analytical | 259.223 | 259.523 | 259.373 | 0 | Polynomial | 0 |  | 259.336 | 259.341 | 259.41 | 259.412 | Range Integration | 259.372 | 259.374 |  |  | True | Y 224.306
Ni 221.648 | ppb | Analytical | 221.498 | 221.798 | 221.648 | 0 | Polynomial | 0 |  | 221.613 | 221.618 | 221.686 | 221.688 | Range Integration | 221.647 | 221.649 |  |  | True | Y 224.306
Ni 231.604 | ppb | Analytical | 231.454 | 231.754 | 231.604 | 0 | Polynomial | 0 |  | 231.578 | 231.579 | 231.633 | 231.635 | Range Integration | 231.603 | 231.605 |  |  | True | Y 224.306
Ni 232.003 | ppb | Analytical | 231.854 | 232.154 | 232.004 | 0.001 | Polynomial | 0 |  | 231.96 | 231.962 | 232.029 | 232.029 | Range Integration | 232.003 | 232.005 |  |  | True | Y 224.306
Pb 220.353 | ppb | Analytical | 220.203 | 220.503 | 220.353 | 0 | Polynomial | 0 |  | 220.32 | 220.323 | 220.378 | 220.381 | Range Integration | 220.352 | 220.354 |  |  | True | Y 224.306
Pb 283.305 | ppb | Analytical | 283.159 | 283.459 | 283.309 | 0.004 | Polynomial | 0 |  | 283.283 | 283.284 | 283.33 | 283.331 | Range Integration | 283.308 | 283.31 |  |  | True | Y 224.306
Se 196.090 | ppb | Analytical | 195.941 | 196.241 | 196.091 | 0.001 | Polynomial | 0 |  | 196.05 | 196.052 |  |  | Range Integration | 196.09 | 196.092 |  |  | True | Y 224.306
Se 204.050 | ppb | Analytical | 203.901 | 204.201 | 204.051 | 0.001 | Polynomial | 0 |  | 204.028 | 204.029 |  |  | Range Integration | 204.05 | 204.052 |  |  | True | Y 224.306
Sr 407.771 | ppb | Analytical | 407.624 | 407.924 | 407.774 | 0.003 | Polynomial | 0 |  | 407.71 | 407.713 | 407.859 | 407.862 | Range Integration | 407.773 | 407.775 |  |  | True | Y 224.306
Sr 421.552 | ppb | Analytical | 421.404 | 421.704 | 421.554 | 0.002 | Polynomial | 0 |  | 421.474 | 421.477 | 421.621 | 421.624 | Range Integration | 421.553 | 421.555 |  |  | True | Y 224.306
Zn 213.856 | ppb | Analytical | 213.708 | 214.008 | 213.858 | 0.002 | Polynomial | 0 |  | 213.83 | 213.831 | 213.891 | 213.893 | Range Integration | 213.857 | 213.859 |  |  | True | Y 224.306
Zn 206.200 | ppb | Analytical | 206.045 | 206.345 | 206.195 | -0.005 | Polynomial | 1 |  | 206.171 | 206.172 | 206.219 | 206.219 | Range Integration | 206.194 | 206.196 |  |  | True | Y 224.306

## Method Regression Descriptions
Line | Display Unit | Regression Model | Degree | Min. Corr. | Min. Concentration | Max. Concentration | Min. Conc. Lock | Max. Conc. Lock | Predefined LOD
Na 589.592 | ppm | Polynomial | 2 | 0.996 | 0 | 5 | False | True
Na 588.995 | ppm | Polynomial | 2 | 0.996 | 0 | 5 | True | True
Na 589.592 R | ppm | Polynomial | 1 | 0.996 | 0 | 500 | True | True
Na 330.298 R | ppm | Polynomial | 1 | 0.996 | 0 | 500 | True | True
K 766.491 | ppm | Polynomial | 2 | 0.996 | 0 | 10 | False | True
K 769.896 | ppm | Polynomial | 2 | 0.996 | 0 | 6 | False | False
K 766.491 R | ppm | Polynomial | 1 | 0.996 | 5 | 240 | True | False
K 769.896 R | ppm | Polynomial | 1 | 0.996 | 5 | 200 | True | True
Mg 285.213 | ppm | Polynomial | 2 | 0.996 | 0 | 60 | False | False
Mg 279.079 | ppm | Polynomial | 1 | 0.996 | 0 | 60 | False | False
Mg 279.079 R | ppm | Polynomial | 1 | 0.996 | 0 | 240 | False | False
Mg 285.213 R | ppm | Polynomial | 1 | 0.996 | 0 | 240 | False | False
Ca 393.366 | ppm | Polynomial | 2 | 0.996 | 0 | 1.2 | False | False
Ca 396.847 | ppm | Polynomial | 2 | 0.996 | 0 | 6 | False | False
Ca 317.933 | ppm | Polynomial | 1 | 0.996 | 0 | 240 | False | False
Ca 422.673 | ppm | Polynomial | 1 | 0.996 | 0 | 240 | False | False
Al 167.078 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Al 396.152 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
As 189.042 | ppb | Polynomial | 1 | 0.996 | 0 | 6000 | False | False
As 193.759 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
B 249.773 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
B 249.677 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Ba 233.527 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Ba 455.404 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Be 234.861 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Be 313.042 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Be 313.107 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Cd 214.438 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Cd 226.502 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Co 228.616 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Co 238.892 | ppb | Polynomial | 1 | 0.996 | 0 | 6000 | False | False
Cr 267.716 | ppb | Polynomial | 1 | 0.996 | 0 | 6000 | False | False
Cr 283.563 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Cu 324.754 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Cu 224.700 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Cu 327.396 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Fe 238.204 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Fe 259.941 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Li 670.780 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Li 670.780 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Mn 257.611 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Mn 259.373 | ppb | Polynomial | 1 | 0.996 | 0 | 6000 | False | False
Ni 221.648 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Ni 231.604 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Ni 232.003 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Pb 220.353 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Pb 283.305 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Se 196.090 | ppb | Polynomial | 1 | 0.996 | 0 | 6000 | False | False
Se 204.050 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Sr 407.771 | ppb | Polynomial | 1 | 0.996 | 0 | 1200 | False | False
Sr 421.552 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False
Zn 213.856 | ppb | Polynomial | 1 | 0.996 | 0 | 6000 | False | False
Zn 206.200 | ppb | Polynomial | 1 | 0.996 | 0 | 12000 | False | False

## Regression Properties
Method | Mjr_Metals_20241201
Version | 20.1.19.1.47.1.61.1.17.1.2.1.5.1.1
Creation Date | 15/03/2023 13:42:15
Regression Date | 06/07/2026 13:38:33

## Regression Line Coeff
Name | Unit | LOD | BEC | Std. Error | Regression Error | Min. Correlation | Correlation | Calib. Range | Coeff. A0 | Coeff. A1 | Coeff. A2 | Sensitivity | Background | Saturation
Na 589.592 | ppm | 0.0012629 | 0.041943 | 0 | None | 0.996 | 1 | 0.000000 - 5.00 | -0.0207 | 0.24939 | -0.001778
Na 588.995 | ppm | 0.00038881 | 0.053703 | 0 | None | 0.996 | 1 | 0.000000 - 5.00 | -0.02478 | 0.1247 | -0.00033024
Na 589.592 R | ppm | 0.37499 | 3.0991 | 0.42197 | None | 0.996 | 0.99999 | 0.000000 - 500 | -0.29195 | 13.8557
Na 330.298 R | ppm | 55.7767 | 422.1756 | 1.3403 | None | 0.996 | 0.99992 | 0.000000 - 500 | 1.9588 | 2080.8983
K 766.491 | ppm | 0.0018175 | 0.09879 | 0.011835 | None | 0.996 | 0.99999 | 0.000000 - 10.00 | -0.018447 | 0.66232 | -0.015231
K 769.896 | ppm | 0.0017294 | 0.12897 | 0.0090595 | None | 0.996 | 0.99999 | 0.000000 - 6.00 | -0.0078945 | 1.995 | -0.14128
K 766.491 R | ppm | 0.97995 | 10.0114 | 0.99495 | None | 0.996 | 0.99995 | 5.00 - 240 | -0.29472 | 48.3455
K 769.896 R | ppm | 1.5844 | 15.5254 | 2.1759 | None | 0.996 | 0.99969 | 5.00 - 200 | 1.4031 | 162.3675
Mg 285.213 | ppm | 0.0010296 | 0.014247 | 0.0081428 | None | 0.996 | 1 | 0.000000 - 60.0 | -0.0013268 | 0.15115 | 7.6039e-05
Mg 279.079 | ppm | 0.023233 | 0.26456 | 0.26315 | None | 0.996 | 0.99994 | 0.000000 - 60.0 | -0.00040529 | 3.9169
Mg 279.079 R | ppm | 0.063562 | 1.4802 | 0.88074 | None | 0.996 | 0.99991 | 0.000000 - 240 | 0.12496 | 26.3231
Mg 285.213 R | ppm | 0.0053697 | 0.092903 | 2.3569 | None | 0.996 | 0.99959 | 0.000000 - 240 | -0.90761 | 1.4865
Ca 393.366 | ppm | 7.2065e-05 | 0.012309 | 0.0057945 | None | 0.996 | 0.99998 | 0.000000 - 1.20 | -0.010866 | 0.0034181 | 1.2731e-06
Ca 396.847 | ppm | 0.00012305 | 0.015609 | 0.010936 | None | 0.996 | 0.99999 | 0.000000 - 6.00 | -0.011843 | 0.0071683 | 1.6448e-06
Ca 317.933 | ppm | 0.065137 | 0.5446 | 0.62264 | None | 0.996 | 0.99998 | 0.000000 - 240 | -0.041794 | 2.2836
Ca 422.673 | ppm | 0.18648 | 1.6299 | 0.20996 | None | 0.996 | 1 | 0.000000 - 240 | 0.046883 | 7.4101
Al 167.078 | ppb | 0.19176 | 2.0778 | 15.5205 | None | 0.996 | 0.99967 | 0.000000 - 1200 | -1.0349 | 138.4545
Al 396.152 | ppb | 1.7859 | 129.1428 | 27.3343 | None | 0.996 | 0.99998 | 0.000000 - 12000 | -9.383 | 609.574
As 189.042 | ppb | 10.4053 | 107.5072 | 9.0393 | None | 0.996 | 0.99999 | 0.000000 - 6000 | 3.2673 | 10168.8506
As 193.759 | ppb | 14.3845 | 146.0231 | 26.4879 | None | 0.996 | 0.99998 | 0.000000 - 12000 | 13.1496 | 13677.0729
B 249.773 | ppb | 1.6018 | 29.4471 | 13.2002 | None | 0.996 | 0.99999 | 0.000000 - 12000 | -1.7358 | 921.8873
B 249.677 | ppb | 4.6063 | 53.5359 | 25.9803 | None | 0.996 | 0.99998 | 0.000000 - 12000 | -3.4908 | 1665.2924
Ba 233.527 | ppb | 1.1857 | 10.2722 | 15.2862 | None | 0.996 | 0.99999 | 0.000000 - 12000 | 0.21498 | 422.2884
Ba 455.404 | ppb | 0.07879 | 4.1138 | 62.8146 | None | 0.996 | 0.99989 | 0.000000 - 12000 | -0.10334 | 21.6233
Be 234.861 | ppb | 0.18438 | 1.88 | 23.1104 | None | 0.996 | 0.99998 | 0.000000 - 12000 | 0.054279 | 77.0945
Be 313.042 | ppb | 0.094716 | 2.5578 | 3.6353 | None | 0.996 | 0.99999 | 0.000000 - 1200 | -0.83242 | 9.6718
Be 313.107 | ppb | 0.084361 | 3.3212 | 14.8214 | None | 0.996 | 0.99999 | 0.000000 - 12000 | 0.51863 | 19.1888
Cd 214.438 | ppb | 0.5049 | 6.0846 | 5.2691 | None | 0.996 | 0.99995 | 0.000000 - 1200 | -0.29373 | 322.0906
Cd 226.502 | ppb | 0.82791 | 8.9867 | 24.4207 | None | 0.996 | 0.99998 | 0.000000 - 12000 | 0.05024 | 403.4717
Co 228.616 | ppb | 2.121 | 20.8237 | 13.8078 | None | 0.996 | 0.99999 | 0.000000 - 12000 | 0.39896 | 933.6194
Co 238.892 | ppb | 2.3553 | 26.4932 | 7.4413 | None | 0.996 | 0.99999 | 0.000000 - 6000 | -0.071086 | 885.9837
Cr 267.716 | ppb | 2.5204 | 24.7468 | 0.95514 | None | 0.996 | 1 | 0.000000 - 6000 | 0.050146 | 816.4426
Cr 283.563 | ppb | 1.8558 | 27.3562 | 12.7128 | None | 0.996 | 1 | 0.000000 - 12000 | -5.5552 | 300.2283
Cu 324.754 | ppb | 0.7659 | 50.7147 | 28.6355 | None | 0.996 | 0.99999 | 0.000000 - 12000 | -9.3999 | 235.2938
Cu 224.700 | ppb | 4.0276 | 53.8933 | 0.99613 | None | 0.996 | 1 | 0.000000 - 1200 | 0.71772 | 2184.7959
Cu 327.396 | ppb | 0.2972 | 81.4877 | 2.212 | None | 0.996 | 0.99999 | 0.000000 - 1200 | 1.2813 | 492.3976
Fe 238.204 | ppb | 1.4204 | 15.4529 | 8.6126 | None | 0.996 | 0.99989 | 0.000000 - 1200 | -0.12937 | 516.615
Fe 259.941 | ppb | 1.725 | 19.3065 | 13.5761 | None | 0.996 | 0.99999 | 0.000000 - 12000 | -6.6023 | 562.3444
Li 670.780 | ppb | 0.024305 | 2.7331 | 8.8386 | None | 0.996 | 1 | 0.000000 - 12000 | -0.078575 | 18.5641
Li 670.780 | ppb | 0.024278 | 2.7331 | 8.8506 | None | 0.996 | 1 | 0.000000 - 12000 | -0.109 | 18.5618
Mn 257.611 | ppb | 0.32037 | 3.8236 | 39.7351 | None | 0.996 | 0.99995 | 0.000000 - 12000 | -0.11458 | 100.9311
Mn 259.373 | ppb | 0.40949 | 4.7886 | 6.9208 | None | 0.996 | 0.99999 | 0.000000 - 6000 | -0.016699 | 134.9267
Ni 221.648 | ppb | 1.9413 | 23.6875 | 5.6239 | None | 0.996 | 0.99994 | 0.000000 - 1200 | -0.24457 | 899.4288
Ni 231.604 | ppb | 3.7141 | 37.2571 | 11.4685 | None | 0.996 | 1 | 0.000000 - 12000 | -0.39478 | 1541.076
Ni 232.003 | ppb | 9.8603 | 98.9352 | 10.0707 | None | 0.996 | 1 | 0.000000 - 12000 | 0.70588 | 4131.0094
Pb 220.353 | ppb | 9.7751 | 132.1568 | 1.3245 | None | 0.996 | 1 | 0.000000 - 1200 | 1.7508 | 5300.7678
Pb 283.305 | ppb | 62.7018 | 700.0646 | 86.7465 | None | 0.996 | 0.99988 | 0.000000 - 12000 | 61.6172 | 9702.2716
Se 196.090 | ppb | 13.6055 | 171.355 | 15.7288 | None | 0.996 | 0.99997 | 0.000000 - 6000 | 10.6653 | 14032.3091
Se 204.050 | ppb | 31.1826 | 410.5784 | 25.3666 | None | 0.996 | 0.99999 | 0.000000 - 12000 | 24.1882 | 24525.7297
Sr 407.771 | ppb | 0.050221 | 1.2983 | 4.6414 | None | 0.996 | 0.99998 | 0.000000 - 1200 | -0.028657 | 6.059
Sr 421.552 | ppb | 0.021026 | 2.8552 | 107.3425 | None | 0.996 | 0.99981 | 0.000000 - 12000 | -61.45 | 18.3084
Zn 213.856 | ppb | 0.84242 | 8.7785 | 2.2056 | None | 0.996 | 1 | 0.000000 - 6000 | -0.76181 | 448.5012
Zn 206.200 | ppb | 0.9296 | 9.6193 | 1.9902 | None | 0.996 | 1 | 0.000000 - 12000 | -1.7953 | 689.3852

## Regression Calibration Points
Na 589.592
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
Blank | Used | 02/07/2026 08:41:04 | 0.08305 | 0.0020863 | 50 | 0 | 0 | -3.6628e-18 | -3.6628e-18 | 100
0.01 | Not used | 02/07/2026 08:45:30 | 0.1436 | 0.0028376 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.33913 | 0.0062398 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 0.54533 | 0.014476 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 2.1928 | 0.070133 | 1 | 0.5 | 0.5
1 | Used | 02/07/2026 09:03:17 | 4.2197 | 0.077984 | 1 | 1 | 1 | 1 | 1 | 6.1107e-13
5 | Used | 02/07/2026 09:07:45 | 24.3641 | 0.15494 | 1 | 5 | 5 | 5 | 5 | -2.8422e-14
10 | Not used | 02/07/2026 09:12:14 | 51.2513 | 0.57376 | 1 | 10 | 10
50 | Not used | 02/07/2026 09:16:42 | 247.887 | 4.3461 | 1 | 50 | 50
100 | Not used | 02/07/2026 09:21:11 | 472.1404 | 3.3415 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 678.8668 | 3.9776 | 1 | 200 | 200
Na 588.995
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
Blank | Used | 02/07/2026 08:41:04 | 0.19882 | 0.0077894 | 50 | 0 | 0 | 3.2614e-18 | 3.2614e-18 | 100
0.01 | Not used | 02/07/2026 08:45:30 | 0.3106 | 0.012003 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.69368 | 0.014083 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 1.0878 | 0.022153 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 4.3485 | 0.14707 | 1 | 0.5 | 0.5
1 | Used | 02/07/2026 09:03:17 | 8.4049 | 0.35096 | 1 | 1 | 1 | 1 | 1 | -5.1159e-13
5 | Used | 02/07/2026 09:07:45 | 45.8651 | 1.2354 | 1 | 5 | 5 | 5 | 5 | 0
10 | Not used | 02/07/2026 09:12:14 | 93.7682 | 0.65388 | 1 | 10 | 10
50 | Not used | 02/07/2026 09:16:42 | 381.8389 | 4.9458 | 1 | 50 | 50
100 | Not used | 02/07/2026 09:21:11 | 627.5385 | 3.4087 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 640.4419 | 4.0174 | 1 | 200 | 200
Na 589.592 R
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.041212 | 0.0023009 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.043 | 0.0020143 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 0.044515 | 0.0024243 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 0.071346 | 0.0038004 | 1 | 0.5 | 0.5
1 | Not used | 02/07/2026 09:03:17 | 0.1057 | 0.0040298 | 1 | 1 | 1
5 | Used | 02/07/2026 09:07:45 | 0.39705 | 0.00060816 | 1 | 5 | 5 | 5.2094 | 5.2094 | 4.0197
10 | Used | 02/07/2026 09:12:14 | 0.753 | 0.010465 | 1 | 10 | 10 | 10.1414 | 10.1414 | 1.3941
50 | Used | 02/07/2026 09:16:42 | 3.5797 | 0.0024492 | 1 | 50 | 50 | 49.3077 | 49.3077 | -1.404
100 | Used | 02/07/2026 09:21:11 | 7.2253 | 0.038903 | 1 | 100 | 100 | 99.8195 | 99.8195 | -0.18079
200 | Used | 02/07/2026 09:25:39 | 14.4735 | 0.12897 | 1 | 200 | 200 | 200.2474 | 200.2474 | 0.12357
Blank | Used | 02/07/2026 08:41:04 | 0.040884 | 0.0017318 | 1 | 0 | 0 | 0.27452 | 0.27452 | 100
Na 330.298 R
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.00021289 | 0.00045857 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.00016181 | 0.0006936 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 0.00090359 | 0.0011432 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 0.002632 | 0.00046393 | 1 | 0.5 | 0.5
1 | Not used | 02/07/2026 09:03:17 | 0.0025323 | 0.00095292 | 1 | 1 | 1
5 | Not used | 02/07/2026 09:07:45 | 0.015508 | 0.0012482 | 1 | 5 | 5
10 | Not used | 02/07/2026 09:12:14 | 0.027545 | 0.00074926 | 1 | 10 | 10
50 | Used | 02/07/2026 09:16:42 | 0.023432 | 0.00075396 | 1 | 50 | 50 | 50.7188 | 50.7188 | 1.4172
100 | Used | 02/07/2026 09:21:11 | 0.047662 | 0.0009426 | 1 | 100 | 100 | 101.1393 | 101.1393 | 1.1264
200 | Used | 02/07/2026 09:25:39 | 0.094802 | 0.0011062 | 1 | 200 | 200 | 199.2327 | 199.2327 | -0.38512
Blank | Used | 02/07/2026 08:41:04 | -0.0014655 | 0.0006179 | 1 | 0 | 0 | -1.0907 | -1.0907 | 100
K 766.491
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.043848 | 0.00070297 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.11304 | 0.0017864 | 1 | 0.05 | 0.05
0.1 | Used | 02/07/2026 08:54:23 | 0.18607 | 0.0031405 | 1 | 0.1 | 0.1 | 0.10426 | 0.10426 | 4.088
0.5 | Used | 02/07/2026 08:58:50 | 0.7746 | 0.020779 | 1 | 0.5 | 0.5 | 0.48545 | 0.48545 | -2.9982
1 | Used | 02/07/2026 09:03:17 | 1.6078 | 0.00972 | 1 | 1 | 1 | 1.0071 | 1.0071 | 0.70299
5 | Used | 02/07/2026 09:07:45 | 9.7738 | 0.025914 | 1 | 5 | 5 | 4.9999 | 4.9999 | -0.0020349
10 | Not used | 02/07/2026 09:12:14 | 21.1221 | 0.19061 | 1 | 10 | 10
50 | Not used | 02/07/2026 09:16:42 | 96.0333 | 1.5618 | 1 | 50 | 50
100 | Not used | 02/07/2026 09:21:11 | 184.0821 | 2.1682 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 336.9157 | 4.669 | 1 | 200 | 200
Blank | Used | 02/07/2026 08:41:04 | 0.027871 | 0.0010228 | 50 | 0 | 0 | 1.3258e-06 | 1.3258e-06 | 100
K 769.896
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.0075628 | 0.001172 | 1 | 0.01 | 0.01 | 0.0071851 | 0.0071851 | -39.1773
0.05 | Used | 02/07/2026 08:49:56 | 0.029314 | 0.00075465 | 1 | 0.05 | 0.05 | 0.050466 | 0.050466 | 0.9229
0.1 | Used | 02/07/2026 08:54:23 | 0.055031 | 0.0017809 | 1 | 0.1 | 0.1 | 0.10146 | 0.10146 | 1.4424
0.5 | Used | 02/07/2026 08:58:50 | 0.25108 | 0.0067147 | 1 | 0.5 | 0.5 | 0.4841 | 0.4841 | -3.2846
1 | Used | 02/07/2026 09:03:17 | 0.52908 | 0.0067791 | 1 | 1 | 1 | 1.0081 | 1.0081 | 0.80079
5 | Used | 02/07/2026 09:07:45 | 3.2652 | 0.0546 | 1 | 5 | 5 | 4.9999 | 4.9999 | -0.0023677
10 | Not used | 02/07/2026 09:12:14 | 7.1183 | 0.079253 | 1 | 10 | 10
50 | Not used | 02/07/2026 09:16:42 | 35.6407 | 0.28527 | 1 | 50 | 50
100 | Not used | 02/07/2026 09:21:11 | 73.6504 | 0.95139 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 147.1304 | 2.0673 | 1 | 200 | 200
Blank | Used | 02/07/2026 08:41:04 | 0.00396 | 0.00052521 | 50 | 0 | 0 | 3.5329e-06 | 3.5329e-06 | 100
K 766.491 R
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.01513 | 0.00080253 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.016918 | 0.00053671 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 0.018729 | 0.00086723 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 0.025826 | 0.0015292 | 1 | 0.5 | 0.5
1 | Not used | 02/07/2026 09:03:17 | 0.036437 | 0.0018286 | 1 | 1 | 1
5 | Used | 02/07/2026 09:07:45 | 0.11902 | 0.0023158 | 1 | 5 | 5 | 5.4593 | 5.4593 | 8.4127
10 | Used | 02/07/2026 09:12:14 | 0.22006 | 0.0022902 | 1 | 10 | 10 | 10.3444 | 10.3444 | 3.3289
50 | Used | 02/07/2026 09:16:42 | 1.0085 | 0.0030931 | 1 | 50 | 50 | 48.461 | 48.461 | -3.1757
100 | Not used | 02/07/2026 09:21:11 | 2.0001 | 0.0046667 | 1 | 100 | 100
200 | Used | 02/07/2026 09:25:39 | 4.15 | 0.074497 | 1 | 200 | 200 | 200.3412 | 200.3412 | 0.17031
Blank | Used | 02/07/2026 08:41:04 | 0.014249 | 0.0007796 | 1 | 0 | 0 | 0.39417 | 0.39417 | 100
K 769.896 R
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | -0.0016458 | 0.00036177 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | -0.0019512 | 0.00031725 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | -0.001912 | 0.00011384 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 0.0003079 | 0.00040232 | 1 | 0.5 | 0.5
1 | Not used | 02/07/2026 09:03:17 | 0.0034535 | 0.00058527 | 1 | 1 | 1
5 | Used | 02/07/2026 09:07:45 | 0.027486 | 0.00022072 | 1 | 5 | 5 | 5.8659 | 5.8659 | 14.7612
10 | Used | 02/07/2026 09:12:14 | 0.057272 | 0.00080857 | 1 | 10 | 10 | 10.7022 | 10.7022 | 6.5616
50 | Used | 02/07/2026 09:16:42 | 0.29075 | 0.0029522 | 1 | 50 | 50 | 48.6122 | 48.6122 | -2.8549
100 | Used | 02/07/2026 09:21:11 | 0.58702 | 0.00053828 | 1 | 100 | 100 | 96.7155 | 96.7155 | -3.396
200 | Used | 02/07/2026 09:25:39 | 1.2345 | 0.021033 | 1 | 200 | 200 | 201.8378 | 201.8378 | 0.91051
Blank | Used | 02/07/2026 08:41:04 | -0.00084144 | 1.8906e-05 | 1 | 0 | 0 | 1.2664 | 1.2664 | 100
Mg 285.213
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.076001 | 0.0010965 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.35077 | 0.0029474 | 1 | 0.05 | 0.05
0.1 | Used | 02/07/2026 08:54:23 | 0.68771 | 0.0039466 | 1 | 0.1 | 0.1 | 0.10266 | 0.10266 | 2.5891
0.5 | Used | 02/07/2026 08:58:50 | 3.376 | 0.018383 | 1 | 0.5 | 0.5 | 0.50983 | 0.50983 | 1.9275
1 | Used | 02/07/2026 09:03:17 | 6.5673 | 0.071253 | 1 | 1 | 1 | 0.99462 | 0.99462 | -0.54123
5 | Not used | 02/07/2026 09:07:45 | 34.6013 | 0.48253 | 1 | 5 | 5
10 | Not used | 02/07/2026 09:12:14 | 69.3887 | 0.091755 | 1 | 10 | 10
50 | Used | 02/07/2026 09:16:42 | 288.8342 | 2.0501 | 1 | 50 | 50 | 50 | 50 | 2.8509e-06
100 | Not used | 02/07/2026 09:21:11 | 534.8922 | 3.5725 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 694.1121 | 4.1878 | 1 | 200 | 200
Blank | Used | 02/07/2026 08:41:04 | 0.0087589 | 0.00032711 | 50 | 0 | 0 | -2.8412e-06 | -2.8412e-06 | 100
Mg 279.079
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.0028379 | 0.0002647 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.013821 | 0.00037467 | 1 | 0.05 | 0.05
0.1 | Used | 02/07/2026 08:54:23 | 0.028293 | 0.00023799 | 1 | 0.1 | 0.1 | 0.11042 | 0.11042 | 9.4329
0.5 | Used | 02/07/2026 08:58:50 | 0.1352 | 0.00058201 | 1 | 0.5 | 0.5 | 0.52915 | 0.52915 | 5.509
1 | Used | 02/07/2026 09:03:17 | 0.26511 | 0.001178 | 1 | 1 | 1 | 1.038 | 1.038 | 3.6617
5 | Used | 02/07/2026 09:07:45 | 1.3448 | 0.0090544 | 1 | 5 | 5 | 5.2672 | 5.2672 | 5.0735
10 | Used | 02/07/2026 09:12:14 | 2.6818 | 0.01018 | 1 | 10 | 10 | 10.504 | 10.504 | 4.7985
50 | Used | 02/07/2026 09:16:42 | 12.7306 | 0.081391 | 1 | 50 | 50 | 49.8645 | 49.8645 | -0.2718
100 | Not used | 02/07/2026 09:21:11 | 27.7857 | 0.30899 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 56.2444 | 0.42424 | 1 | 200 | 200
Blank | Used | 02/07/2026 08:41:04 | 3.0628e-05 | 0.00013501 | 50 | 0 | 0 | -0.00028532 | -0.00028532 | 100
Mg 279.079 R
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.00019168 | 0.000264 | 1 | 0.01 | 0.01 | 0.13001 | 0.13001 | 92.3083
0.05 | Used | 02/07/2026 08:49:56 | 0.0022841 | 0.00053612 | 1 | 0.05 | 0.05 | 0.18509 | 0.18509 | 72.986
0.1 | Used | 02/07/2026 08:54:23 | 0.0041613 | 0.00060636 | 1 | 0.1 | 0.1 | 0.2345 | 0.2345 | 57.3565
0.5 | Used | 02/07/2026 08:58:50 | 0.019935 | 0.00094228 | 1 | 0.5 | 0.5 | 0.64973 | 0.64973 | 23.0445
1 | Used | 02/07/2026 09:03:17 | 0.039901 | 0.0016723 | 1 | 1 | 1 | 1.1753 | 1.1753 | 14.9136
5 | Used | 02/07/2026 09:07:45 | 0.19867 | 0.0043237 | 1 | 5 | 5 | 5.3546 | 5.3546 | 6.6228
10 | Used | 02/07/2026 09:12:14 | 0.38651 | 0.0071705 | 1 | 10 | 10 | 10.2992 | 10.2992 | 2.9051
50 | Used | 02/07/2026 09:16:42 | 1.8016 | 0.037893 | 1 | 50 | 50 | 47.5496 | 47.5496 | -5.1534
100 | Used | 02/07/2026 09:21:11 | 3.8239 | 0.034461 | 1 | 100 | 100 | 100.7826 | 100.7826 | 0.7765
200 | Used | 02/07/2026 09:25:39 | 7.5993 | 0.066435 | 1 | 200 | 200 | 200.1612 | 200.1612 | 0.080548
Blank | Used | 02/07/2026 08:41:04 | 0.0005023 | 0.0005443 | 1 | 0 | 0 | 0.13819 | 0.13819 | 100
Mg 285.213 R
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.007477 | 0.0003246 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.039699 | 0.00038432 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 0.080468 | 0.00071615 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 0.39287 | 0.010303 | 1 | 0.5 | 0.5
1 | Used | 02/07/2026 09:03:17 | 0.74542 | 0.013016 | 1 | 1 | 1 | 0.20045 | 0.20045 | -398.872
5 | Used | 02/07/2026 09:07:45 | 3.8138 | 0.025484 | 1 | 5 | 5 | 4.7616 | 4.7616 | -5.0076
10 | Used | 02/07/2026 09:12:14 | 7.5157 | 0.085408 | 1 | 10 | 10 | 10.2644 | 10.2644 | 2.5759
50 | Used | 02/07/2026 09:16:42 | 33.8427 | 0.45004 | 1 | 50 | 50 | 49.3994 | 49.3994 | -1.2158
100 | Used | 02/07/2026 09:21:11 | 70.9426 | 0.74405 | 1 | 100 | 100 | 104.5484 | 104.5484 | 4.3505
200 | Used | 02/07/2026 09:25:39 | 133.6305 | 1.1203 | 1 | 200 | 200 | 197.7338 | 197.7338 | -1.1461
Blank | Used | 02/07/2026 08:41:04 | -0.00029993 | 0.00081613 | 1 | 0 | 0 | -0.90805 | -0.90805 | 100
Ca 393.366
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
Blank | Used | 02/07/2026 08:41:04 | 3.1749 | 0.0206 | 50 | 0 | 0 | -1.4776e-06 | -1.4776e-06 | 100
0.01 | Not used | 02/07/2026 08:45:30 | 7.1418 | 0.13588 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 20.1768 | 0.15551 | 1 | 0.05 | 0.05
0.1 | Used | 02/07/2026 08:54:23 | 33.58 | 0.17774 | 1 | 0.1 | 0.1 | 0.10535 | 0.10535 | 5.0762
0.5 | Used | 02/07/2026 08:58:50 | 141.3811 | 0.89653 | 1 | 0.5 | 0.5 | 0.49783 | 0.49783 | -0.43601
1 | Used | 02/07/2026 09:03:17 | 268.9527 | 5.4222 | 1 | 1 | 1 | 1.0005 | 1.0005 | 0.051669
5 | Not used | 02/07/2026 09:07:45 | 680.7409 | 4.3561 | 1 | 5 | 5
10 | Not used | 02/07/2026 09:12:14 | 664.8223 | 5.3865 | 1 | 10 | 10
50 | Not used | 02/07/2026 09:16:42 | 656.9 | 6.7119 | 1 | 50 | 50
100 | Not used | 02/07/2026 09:21:11 | 656.5876 | 2.8477 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 644.4826 | 3.6594 | 1 | 200 | 200
Ca 396.847
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 3.7037 | 0.057802 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 10.2932 | 0.080497 | 1 | 0.05 | 0.05
0.1 | Used | 02/07/2026 08:54:23 | 17.0013 | 0.088538 | 1 | 0.1 | 0.1 | 0.1105 | 0.1105 | 9.5038
0.5 | Used | 02/07/2026 08:58:50 | 71.4947 | 0.44881 | 1 | 0.5 | 0.5 | 0.50906 | 0.50906 | 1.779
1 | Used | 02/07/2026 09:03:17 | 135.9597 | 2.8138 | 1 | 1 | 1 | 0.99315 | 0.99315 | -0.68927
5 | Used | 02/07/2026 09:07:45 | 612.9839 | 9.9354 | 1 | 5 | 5 | 5.0002 | 5.0002 | 0.0041115
10 | Not used | 02/07/2026 09:12:14 | 685.2247 | 5.4535 | 1 | 10 | 10
50 | Not used | 02/07/2026 09:16:42 | 662.7152 | 6.7284 | 1 | 50 | 50
100 | Not used | 02/07/2026 09:21:11 | 674.1363 | 2.851 | 1 | 100 | 100
200 | Not used | 02/07/2026 09:25:39 | 666.5285 | 3.7811 | 1 | 200 | 200
Blank | Used | 02/07/2026 08:41:04 | 1.6508 | 0.0094291 | 50 | 0 | 0 | -5.1673e-06 | -5.1673e-06 | 100
Ca 317.933
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.024793 | 0.00016376 | 1 | 0.01 | 0.01
0.05 | Not used | 02/07/2026 08:49:56 | 0.048021 | 0.0008288 | 1 | 0.05 | 0.05
0.1 | Not used | 02/07/2026 08:54:23 | 0.071783 | 0.00054737 | 1 | 0.1 | 0.1
0.5 | Not used | 02/07/2026 08:58:50 | 0.25707 | 0.0033972 | 1 | 0.5 | 0.5
1 | Used | 02/07/2026 09:03:17 | 0.4752 | 0.0015794 | 1 | 1 | 1 | 1.0434 | 1.0434 | 4.1588
5 | Used | 02/07/2026 09:07:45 | 2.2895 | 0.0096651 | 1 | 5 | 5 | 5.1865 | 5.1865 | 3.5966
10 | Used | 02/07/2026 09:12:14 | 4.5155 | 0.040644 | 1 | 10 | 10 | 10.27 | 10.27 | 2.6291
50 | Used | 02/07/2026 09:16:42 | 22.0283 | 0.075323 | 1 | 50 | 50 | 50.2628 | 50.2628 | 0.52288
100 | Used | 02/07/2026 09:21:11 | 44.3103 | 0.62513 | 1 | 100 | 100 | 101.1466 | 101.1466 | 1.1336
200 | Used | 02/07/2026 09:25:39 | 87.3059 | 0.30194 | 1 | 200 | 200 | 199.3329 | 199.3329 | -0.33466
Blank | Used | 02/07/2026 08:41:04 | 0.018084 | 0.0010549 | 50 | 0 | 0 | -0.00049691 | -0.00049691 | 100
Ca 422.673
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppm] | Prep. Corr. Cert. Conc. [ppm] | Calc. Conc. [ppm] | Calc. Conc. Uncorr. [ppm] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.00030734 | 0.00054862 | 1 | 0.01 | 0.01
0.05 | Used | 02/07/2026 08:49:56 | 0.0067302 | 0.00022275 | 1 | 0.05 | 0.05 | 0.096754 | 0.096754 | 48.3227
0.1 | Used | 02/07/2026 08:54:23 | 0.014122 | 0.00079277 | 1 | 0.1 | 0.1 | 0.15153 | 0.15153 | 34.0066
0.5 | Used | 02/07/2026 08:58:50 | 0.070535 | 0.00099972 | 1 | 0.5 | 0.5 | 0.56955 | 0.56955 | 12.2117
1 | Used | 02/07/2026 09:03:17 | 0.13811 | 0.001868 | 1 | 1 | 1 | 1.0703 | 1.0703 | 6.5692
5 | Used | 02/07/2026 09:07:45 | 0.68179 | 0.0069057 | 1 | 5 | 5 | 5.0991 | 5.0991 | 1.9427
10 | Used | 02/07/2026 09:12:14 | 1.3482 | 0.018575 | 1 | 10 | 10 | 10.0373 | 10.0373 | 0.37132
50 | Used | 02/07/2026 09:16:42 | 6.6663 | 0.079397 | 1 | 50 | 50 | 49.4452 | 49.4452 | -1.122
100 | Used | 02/07/2026 09:21:11 | 13.4893 | 0.053433 | 1 | 100 | 100 | 100.0039 | 100.0039 | 0.0039386
200 | Used | 02/07/2026 09:25:39 | 27.0013 | 0.20403 | 1 | 200 | 200 | 200.1301 | 200.1301 | 0.064987
Blank | Used | 02/07/2026 08:41:04 | -7.8604e-05 | 0.00046827 | 1 | 0 | 0 | 0.0463 | 0.0463 | 100
Al 167.078
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.17589 | 0.0029016 | 1 | 10 | 10 | 23.3177 | 23.3177 | 57.1141
0.05 | Used | 02/07/2026 08:49:56 | 0.43451 | 0.0070809 | 1 | 50 | 50 | 59.1251 | 59.1251 | 15.4336
0.1 | Used | 02/07/2026 08:54:23 | 0.79732 | 0.0058145 | 1 | 100 | 100 | 109.3579 | 109.3579 | 8.5571
0.5 | Used | 02/07/2026 08:58:50 | 3.7712 | 0.024455 | 1 | 500 | 500 | 521.1021 | 521.1021 | 4.0495
1 | Used | 02/07/2026 09:03:17 | 7.1359 | 0.048494 | 1 | 1000 | 1000 | 986.9596 | 986.9596 | -1.3213
5 | Not used | 02/07/2026 09:07:45 | 33.9069 | 0.19917 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 61.9321 | 0.49999 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.0073596 | 0.00027854 | 50 | 0 | 0 | -0.015945 | -0.015945 | 100
Al 396.152
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.039933 | 0.0021998 | 1 | 10 | 10
0.05 | Used | 02/07/2026 08:49:56 | 0.10042 | 0.0034342 | 1 | 50 | 50 | 51.8314 | 51.8314 | 3.5333
0.1 | Used | 02/07/2026 08:54:23 | 0.17825 | 0.0052377 | 1 | 100 | 100 | 99.271 | 99.271 | -0.73433
0.5 | Used | 02/07/2026 08:58:50 | 0.83325 | 0.025565 | 1 | 500 | 500 | 498.5437 | 498.5437 | -0.29212
1 | Used | 02/07/2026 09:03:17 | 1.6243 | 0.01465 | 1 | 1000 | 1000 | 980.7301 | 980.7301 | -1.9649
5 | Used | 02/07/2026 09:07:45 | 8.3031 | 0.043664 | 1 | 5000 | 5000 | 5051.9728 | 5051.9728 | 1.0288
10 | Used | 02/07/2026 09:12:14 | 16.3803 | 0.17267 | 1 | 10000 | 10000 | 9975.638 | 9975.638 | -0.24422
Blank | Used | 02/07/2026 08:41:04 | 0.0022903 | 0.0010168 | 1 | 0 | 0 | -7.9869 | -7.9869 | 100
As 189.042
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.00070444 | 1.986e-05 | 1 | 10 | 10 | 10.4306 | 10.4306 | 4.1284
0.05 | Used | 02/07/2026 08:49:56 | 0.0046469 | 0.0002186 | 1 | 50 | 50 | 50.5205 | 50.5205 | 1.0303
0.1 | Used | 02/07/2026 08:54:23 | 0.0094788 | 0.00015254 | 1 | 100 | 100 | 99.6561 | 99.6561 | -0.34513
0.5 | Used | 02/07/2026 08:58:50 | 0.047381 | 0.00032803 | 1 | 500 | 500 | 485.0802 | 485.0802 | -3.0757
1 | Used | 02/07/2026 09:03:17 | 0.096739 | 0.00037074 | 1 | 1000 | 1000 | 986.9899 | 986.9899 | -1.3182
5 | Used | 02/07/2026 09:07:45 | 0.49177 | 0.0010391 | 1 | 5000 | 5000 | 5004.0131 | 5004.0131 | 0.080197
10 | Not used | 02/07/2026 09:12:14 | 1.0038 | 0.0058953 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | -0.00032039 | 9.3294e-05 | 50 | 0 | 0 | 0.0093239 | 0.0093239 | 100
As 193.759
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.00073679 | 7.7961e-05 | 1 | 10 | 10 | 23.2268 | 23.2268 | 56.9462
0.05 | Used | 02/07/2026 08:49:56 | 0.0037056 | 0.0001528 | 1 | 50 | 50 | 63.8308 | 63.8308 | 21.6679
0.1 | Used | 02/07/2026 08:54:23 | 0.0071799 | 4.7738e-05 | 1 | 100 | 100 | 111.3497 | 111.3497 | 10.1929
0.5 | Used | 02/07/2026 08:58:50 | 0.03501 | 2.8601e-05 | 1 | 500 | 500 | 491.9854 | 491.9854 | -1.629
1 | Used | 02/07/2026 09:03:17 | 0.070437 | 0.00075146 | 1 | 1000 | 1000 | 976.5164 | 976.5164 | -2.4048
5 | Used | 02/07/2026 09:07:45 | 0.36116 | 0.0019729 | 1 | 5000 | 5000 | 4952.8227 | 4952.8227 | -0.95253
10 | Used | 02/07/2026 09:12:14 | 0.73207 | 0.0027002 | 1 | 10000 | 10000 | 10025.7209 | 10025.7209 | 0.25655
Blank | Used | 02/07/2026 08:41:04 | 0.0001022 | 1.8632e-05 | 1 | 0 | 0 | 14.5474 | 14.5474 | 100
B 249.773
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.022371 | 0.0012931 | 1 | 10 | 10 | 18.8877 | 18.8877 | 47.0554
0.05 | Used | 02/07/2026 08:49:56 | 0.067846 | 0.00061688 | 1 | 50 | 50 | 60.8103 | 60.8103 | 17.7772
0.1 | Used | 02/07/2026 08:54:23 | 0.12702 | 0.0011461 | 1 | 100 | 100 | 115.361 | 115.361 | 13.3156
0.5 | Used | 02/07/2026 08:58:50 | 0.54008 | 0.0021233 | 1 | 500 | 500 | 496.1553 | 496.1553 | -0.77489
1 | Used | 02/07/2026 09:03:17 | 1.0682 | 0.0042647 | 1 | 1000 | 1000 | 983.0265 | 983.0265 | -1.7267
5 | Used | 02/07/2026 09:07:45 | 5.4092 | 0.039846 | 1 | 5000 | 5000 | 4984.9099 | 4984.9099 | -0.30272
10 | Used | 02/07/2026 09:12:14 | 10.8591 | 0.044924 | 1 | 10000 | 10000 | 10009.1135 | 10009.1135 | 0.091053
Blank | Used | 02/07/2026 08:41:04 | 0.0018793 | 0.00045162 | 50 | 0 | 0 | -0.0033057 | -0.0033057 | 100
B 249.677
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.012015 | 0.0003497 | 1 | 10 | 10 | 16.5183 | 16.5183 | 39.4612
0.05 | Used | 02/07/2026 08:49:56 | 0.035589 | 0.00054939 | 1 | 50 | 50 | 55.776 | 55.776 | 10.3557
0.1 | Used | 02/07/2026 08:54:23 | 0.067731 | 0.00075389 | 1 | 100 | 100 | 109.3006 | 109.3006 | 8.5092
0.5 | Used | 02/07/2026 08:58:50 | 0.29817 | 0.0013364 | 1 | 500 | 500 | 493.0447 | 493.0447 | -1.4107
1 | Used | 02/07/2026 09:03:17 | 0.58935 | 0.0024273 | 1 | 1000 | 1000 | 977.9545 | 977.9545 | -2.2542
5 | Used | 02/07/2026 09:07:45 | 2.974 | 0.012659 | 1 | 5000 | 5000 | 4949.0399 | 4949.0399 | -1.0297
10 | Used | 02/07/2026 09:12:14 | 6.0236 | 0.035287 | 1 | 10000 | 10000 | 10027.4989 | 10027.4989 | 0.27424
Blank | Used | 02/07/2026 08:41:04 | 0.0021036 | 7.3551e-05 | 50 | 0 | 0 | 0.012347 | 0.012347 | 100
Ba 233.527
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.023378 | 0.00093383 | 1 | 10 | 10
0.05 | Used | 02/07/2026 08:49:56 | 0.12267 | 0.0018784 | 1 | 50 | 50 | 52.0188 | 52.0188 | 3.8808
0.1 | Used | 02/07/2026 08:54:23 | 0.24618 | 0.003111 | 1 | 100 | 100 | 104.1744 | 104.1744 | 4.0072
0.5 | Used | 02/07/2026 08:58:50 | 1.1969 | 0.0037684 | 1 | 500 | 500 | 505.6352 | 505.6352 | 1.1145
1 | Used | 02/07/2026 09:03:17 | 2.3446 | 0.0055894 | 1 | 1000 | 1000 | 990.3015 | 990.3015 | -0.97935
5 | Used | 02/07/2026 09:07:45 | 11.9079 | 0.085918 | 1 | 5000 | 5000 | 5028.7847 | 5028.7847 | 0.5724
10 | Used | 02/07/2026 09:12:14 | 23.6471 | 0.067177 | 1 | 10000 | 10000 | 9986.1271 | 9986.1271 | -0.13892
Blank | Used | 02/07/2026 08:41:04 | -0.00052522 | 9.8113e-05 | 50 | 0 | 0 | -0.0068167 | -0.0068167 | 100
Ba 455.404
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.51274 | 0.008199 | 1 | 10 | 10 | 10.9838 | 10.9838 | 8.9569
0.05 | Used | 02/07/2026 08:49:56 | 2.6022 | 0.032217 | 1 | 50 | 50 | 56.164 | 56.164 | 10.975
0.1 | Used | 02/07/2026 08:54:23 | 5.1493 | 0.027894 | 1 | 100 | 100 | 111.2417 | 111.2417 | 10.1057
0.5 | Used | 02/07/2026 08:58:50 | 24.7133 | 0.65498 | 1 | 500 | 500 | 534.2803 | 534.2803 | 6.4162
1 | Used | 02/07/2026 09:03:17 | 48.3334 | 0.90447 | 1 | 1000 | 1000 | 1045.0256 | 1045.0256 | 4.3086
5 | Used | 02/07/2026 09:07:45 | 236.9607 | 4.7655 | 1 | 5000 | 5000 | 5123.7728 | 5123.7728 | 2.4157
10 | Used | 02/07/2026 09:12:14 | 459.2028 | 3.2944 | 1 | 10000 | 10000 | 9929.3845 | 9929.3845 | -0.71118
Blank | Used | 02/07/2026 08:41:04 | 0.0019885 | 0.00078708 | 50 | 0 | 0 | -0.060341 | -0.060341 | 100
Be 234.861
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.12835 | 0.00090733 | 1 | 10 | 10 | 9.9492 | 9.9492 | -0.51025
0.05 | Used | 02/07/2026 08:49:56 | 0.64927 | 0.0033624 | 1 | 50 | 50 | 50.1092 | 50.1092 | 0.21794
0.1 | Used | 02/07/2026 08:54:23 | 1.2942 | 0.0039157 | 1 | 100 | 100 | 99.8311 | 99.8311 | -0.16916
0.5 | Used | 02/07/2026 08:58:50 | 6.3929 | 0.1194 | 1 | 500 | 500 | 492.912 | 492.912 | -1.438
1 | Used | 02/07/2026 09:03:17 | 12.5374 | 0.11579 | 1 | 1000 | 1000 | 966.6213 | 966.6213 | -3.4531
5 | Used | 02/07/2026 09:07:45 | 65.3951 | 0.51427 | 1 | 5000 | 5000 | 5041.6582 | 5041.6582 | 0.82628
10 | Used | 02/07/2026 09:12:14 | 129.4838 | 0.23113 | 1 | 10000 | 10000 | 9982.5439 | 9982.5439 | -0.17487
Blank | Used | 02/07/2026 08:41:04 | -0.00061909 | 8.0701e-05 | 50 | 0 | 0 | 0.00655 | 0.00655 | 100
Be 313.042
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 1.1258 | 0.02102 | 1 | 10 | 10
0.05 | Used | 02/07/2026 08:49:56 | 5.4417 | 0.019927 | 1 | 50 | 50 | 51.7988 | 51.7988 | 3.4727
0.1 | Not used | 02/07/2026 08:54:23 | 10.787 | 0.048482 | 1 | 100 | 100
0.5 | Used | 02/07/2026 08:58:50 | 52.2233 | 0.4072 | 1 | 500 | 500 | 504.2601 | 504.2601 | 0.84481
1 | Used | 02/07/2026 09:03:17 | 103.2474 | 1.8158 | 1 | 1000 | 1000 | 997.7536 | 997.7536 | -0.22515
5 | Not used | 02/07/2026 09:07:45 | 521.5868 | 7.6566 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 692.6667 | 5.8627 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.085909 | 0.0036076 | 50 | 0 | 0 | -0.001525 | -0.001525 | 100
Be 313.107
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.48874 | 0.0060344 | 1 | 10 | 10 | 9.897 | 9.897 | -1.041
0.05 | Used | 02/07/2026 08:49:56 | 2.5922 | 0.024252 | 1 | 50 | 50 | 50.2595 | 50.2595 | 0.51638
0.1 | Used | 02/07/2026 08:54:23 | 5.2541 | 0.05054 | 1 | 100 | 100 | 101.339 | 101.339 | 1.3213
0.5 | Not used | 02/07/2026 08:58:50 | 25.7144 | 0.46392 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 50.9969 | 0.75483 | 1 | 1000 | 1000 | 979.0875 | 979.0875 | -2.1359
5 | Used | 02/07/2026 09:07:45 | 261.7766 | 3.8856 | 1 | 5000 | 5000 | 5023.6965 | 5023.6965 | 0.47169
10 | Used | 02/07/2026 09:12:14 | 520.5953 | 6.168 | 1 | 10000 | 10000 | 9990.1186 | 9990.1186 | -0.098912
Blank | Used | 02/07/2026 08:41:04 | -0.026911 | 0.0021945 | 50 | 0 | 0 | 0.0022408 | 0.0022408 | 100
Cd 214.438
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.032303 | 0.0011616 | 1 | 10 | 10 | 10.1108 | 10.1108 | 1.0959
0.05 | Not used | 02/07/2026 08:49:56 | 0.16086 | 0.0022406 | 1 | 50 | 50
0.1 | Used | 02/07/2026 08:54:23 | 0.32256 | 0.006729 | 1 | 100 | 100 | 103.6001 | 103.6001 | 3.475
0.5 | Used | 02/07/2026 08:58:50 | 1.576 | 0.011732 | 1 | 500 | 500 | 507.3141 | 507.3141 | 1.4417
1 | Used | 02/07/2026 09:03:17 | 3.0929 | 0.024003 | 1 | 1000 | 1000 | 995.8985 | 995.8985 | -0.41184
5 | Not used | 02/07/2026 09:07:45 | 15.7003 | 0.1129 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 30.9177 | 0.38088 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.00090336 | 3.9012e-05 | 50 | 0 | 0 | -0.0027694 | -0.0027694 | 100
Cd 226.502
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.024649 | 0.00072152 | 1 | 10 | 10 | 9.9954 | 9.9954 | -0.046452
0.05 | Used | 02/07/2026 08:49:56 | 0.12796 | 0.0019946 | 1 | 50 | 50 | 51.6781 | 51.6781 | 3.2473
0.1 | Used | 02/07/2026 08:54:23 | 0.25722 | 0.0039586 | 1 | 100 | 100 | 103.8314 | 103.8314 | 3.69
0.5 | Used | 02/07/2026 08:58:50 | 1.2557 | 0.0032008 | 1 | 500 | 500 | 506.6903 | 506.6903 | 1.3204
1 | Not used | 02/07/2026 09:03:17 | 2.4543 | 0.012858 | 1 | 1000 | 1000
5 | Used | 02/07/2026 09:07:45 | 12.5114 | 0.10766 | 1 | 5000 | 5000 | 5048.0541 | 5048.0541 | 0.95193
10 | Used | 02/07/2026 09:12:14 | 24.7235 | 0.2161 | 1 | 10000 | 10000 | 9975.2935 | 9975.2935 | -0.24768
Blank | Used | 02/07/2026 08:41:04 | -0.00015976 | 9.6508e-05 | 50 | 0 | 0 | -0.014217 | -0.014217 | 100
Co 228.616
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.010275 | 0.00039253 | 1 | 10 | 10 | 9.9915 | 9.9915 | -0.085445
0.05 | Used | 02/07/2026 08:49:56 | 0.054088 | 0.0018519 | 1 | 50 | 50 | 50.8963 | 50.8963 | 1.7611
0.1 | Used | 02/07/2026 08:54:23 | 0.11105 | 0.0019027 | 1 | 100 | 100 | 104.0731 | 104.0731 | 3.9137
0.5 | Used | 02/07/2026 08:58:50 | 0.54556 | 0.0014182 | 1 | 500 | 500 | 509.7399 | 509.7399 | 1.9108
1 | Used | 02/07/2026 09:03:17 | 1.0636 | 0.00034034 | 1 | 1000 | 1000 | 993.4275 | 993.4275 | -0.6616
5 | Used | 02/07/2026 09:07:45 | 5.3852 | 0.047243 | 1 | 5000 | 5000 | 5028.1252 | 5028.1252 | 0.55936
10 | Used | 02/07/2026 09:12:14 | 10.6955 | 0.045388 | 1 | 10000 | 10000 | 9985.9481 | 9985.9481 | -0.14072
Blank | Used | 02/07/2026 08:41:04 | -0.00043684 | 7.5048e-05 | 50 | 0 | 0 | -0.0088806 | -0.0088806 | 100
Co 238.892
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.011407 | 0.00036246 | 1 | 10 | 10 | 10.0349 | 10.0349 | 0.34779
0.05 | Used | 02/07/2026 08:49:56 | 0.057236 | 0.0019086 | 1 | 50 | 50 | 50.6393 | 50.6393 | 1.2624
0.1 | Used | 02/07/2026 08:54:23 | 0.11702 | 0.0013298 | 1 | 100 | 100 | 103.6025 | 103.6025 | 3.4772
0.5 | Used | 02/07/2026 08:58:50 | 0.56731 | 0.0019412 | 1 | 500 | 500 | 502.5561 | 502.5561 | 0.50862
1 | Used | 02/07/2026 09:03:17 | 1.1109 | 0.0028905 | 1 | 1000 | 1000 | 984.2114 | 984.2114 | -1.6042
5 | Used | 02/07/2026 09:07:45 | 5.6466 | 0.039926 | 1 | 5000 | 5000 | 5002.7682 | 5002.7682 | 0.055334
10 | Not used | 02/07/2026 09:12:14 | 11.2316 | 0.061173 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 8.3027e-05 | 0.00010274 | 50 | 0 | 0 | 0.0024751 | 0.0024751 | 100
Cr 267.716
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.012281 | 0.00039444 | 1 | 10 | 10 | 10.0767 | 10.0767 | 0.76133
0.05 | Used | 02/07/2026 08:49:56 | 0.062832 | 0.0011291 | 1 | 50 | 50 | 51.3485 | 51.3485 | 2.6262
0.1 | Not used | 02/07/2026 08:54:23 | 0.12759 | 0.0014903 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 0.617 | 0.0015179 | 1 | 500 | 500
1 | Not used | 02/07/2026 09:03:17 | 1.2054 | 0.0023573 | 1 | 1000 | 1000
5 | Used | 02/07/2026 09:07:45 | 6.1241 | 0.046876 | 1 | 5000 | 5000 | 4999.986 | 4999.986 | -0.00028007
10 | Not used | 02/07/2026 09:12:14 | 12.218 | 0.054707 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | -6.2111e-05 | 6.9842e-05 | 50 | 0 | 0 | -0.00056449 | -0.00056449 | 100
Cr 283.563
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.051885 | 0.0010936 | 1 | 10 | 10
0.05 | Not used | 02/07/2026 08:49:56 | 0.18842 | 0.0018026 | 1 | 50 | 50
0.1 | Not used | 02/07/2026 08:54:23 | 0.36244 | 0.0026791 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 1.6968 | 0.0013818 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 3.301 | 0.0062983 | 1 | 1000 | 1000 | 985.4885 | 985.4885 | -1.4725
5 | Used | 02/07/2026 09:07:45 | 16.7058 | 0.14571 | 1 | 5000 | 5000 | 5009.9925 | 5009.9925 | 0.19945
10 | Used | 02/07/2026 09:12:14 | 33.3146 | 0.14115 | 1 | 10000 | 10000 | 9996.4226 | 9996.4226 | -0.035787
Blank | Used | 02/07/2026 08:41:04 | 0.018514 | 0.00031571 | 50 | 0 | 0 | 0.0032386 | 0.0032386 | 100
Cu 324.754
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.080477 | 0.0021197 | 1 | 10 | 10 | 9.5358 | 9.5358 | -4.868
0.05 | Used | 02/07/2026 08:49:56 | 0.26006 | 0.0024351 | 1 | 50 | 50 | 51.7912 | 51.7912 | 3.4586
0.1 | Not used | 02/07/2026 08:54:23 | 0.48265 | 0.0088138 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 2.1769 | 0.065824 | 1 | 500 | 500
1 | Not used | 02/07/2026 09:03:17 | 4.2518 | 0.11625 | 1 | 1000 | 1000
5 | Used | 02/07/2026 09:07:45 | 21.4779 | 0.18941 | 1 | 5000 | 5000 | 5044.2287 | 5044.2287 | 0.87682
10 | Used | 02/07/2026 09:12:14 | 42.4449 | 0.46259 | 1 | 10000 | 10000 | 9977.6311 | 9977.6311 | -0.22419
Blank | Used | 02/07/2026 08:41:04 | 0.03991 | 0.00057066 | 50 | 0 | 0 | -0.0092748 | -0.0092748 | 100
Cu 224.700
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.0042412 | 0.00032742 | 1 | 10 | 10 | 9.9838 | 9.9838 | -0.16202
0.05 | Used | 02/07/2026 08:49:56 | 0.023201 | 0.00071188 | 1 | 50 | 50 | 51.4068 | 51.4068 | 2.7366
0.1 | Not used | 02/07/2026 08:54:23 | 0.046851 | 0.0012813 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 0.23369 | 0.001447 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 0.45735 | 0.00056501 | 1 | 1000 | 1000 | 999.9278 | 999.9278 | -0.0072168
5 | Not used | 02/07/2026 09:07:45 | 2.3034 | 0.014429 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 4.5875 | 0.022855 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | -0.00032875 | 0.00010889 | 50 | 0 | 0 | -0.00052738 | -0.00052738 | 100
Cu 327.396
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.016426 | 0.00036644 | 1 | 10 | 10 | 9.3693 | 9.3693 | -6.7315
0.05 | Used | 02/07/2026 08:49:56 | 0.10516 | 0.0030196 | 1 | 50 | 50 | 53.06 | 53.06 | 5.7671
0.1 | Not used | 02/07/2026 08:54:23 | 0.21215 | 0.005089 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 1.0261 | 0.03101 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 2.028 | 0.012824 | 1 | 1000 | 1000 | 999.8435 | 999.8435 | -0.015651
5 | Not used | 02/07/2026 09:07:45 | 10.3605 | 0.09267 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 20.5235 | 0.21334 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | -0.002604 | 0.00046319 | 50 | 0 | 0 | -0.00090913 | -0.00090913 | 100
Fe 238.204
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.020013 | 0.00078759 | 1 | 10 | 10
0.05 | Used | 02/07/2026 08:49:56 | 0.10167 | 0.0010177 | 1 | 50 | 50 | 52.3944 | 52.3944 | 4.5699
0.1 | Used | 02/07/2026 08:54:23 | 0.20575 | 0.0021584 | 1 | 100 | 100 | 106.1622 | 106.1622 | 5.8045
0.5 | Used | 02/07/2026 08:58:50 | 0.99045 | 0.0018672 | 1 | 500 | 500 | 511.5527 | 511.5527 | 2.2584
1 | Used | 02/07/2026 09:03:17 | 1.9229 | 0.00090439 | 1 | 1000 | 1000 | 993.2651 | 993.2651 | -0.67806
5 | Not used | 02/07/2026 09:07:45 | 9.7592 | 0.065406 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 19.3219 | 0.10159 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.00024007 | 4.919e-05 | 50 | 0 | 0 | -0.0053498 | -0.0053498 | 100
Fe 259.941
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.018659 | 0.00078101 | 1 | 10 | 10 | 3.8905 | 3.8905 | -157.0381
0.05 | Used | 02/07/2026 08:49:56 | 0.092753 | 0.00049883 | 1 | 50 | 50 | 45.5568 | 45.5568 | -9.7531
0.1 | Used | 02/07/2026 08:54:23 | 0.18877 | 0.0018408 | 1 | 100 | 100 | 99.5521 | 99.5521 | -0.44992
0.5 | Used | 02/07/2026 08:58:50 | 0.90978 | 0.0029843 | 1 | 500 | 500 | 505.0098 | 505.0098 | 0.99203
1 | Not used | 02/07/2026 09:03:17 | 1.7697 | 0.0040898 | 1 | 1000 | 1000
5 | Used | 02/07/2026 09:07:45 | 8.9478 | 0.054601 | 1 | 5000 | 5000 | 5025.1684 | 5025.1684 | 0.50085
10 | Used | 02/07/2026 09:12:14 | 17.7715 | 0.071231 | 1 | 10000 | 10000 | 9987.1059 | 9987.1059 | -0.12911
Blank | Used | 02/07/2026 08:41:04 | 0.00056677 | 6.5985e-05 | 1 | 0 | 0 | -6.2836 | -6.2836 | 100
Li 670.780
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.4904 | 0.0074778 | 1 | 10 | 10 | 9.0252 | 9.0252 | -10.8004
0.05 | Used | 02/07/2026 08:49:56 | 2.5146 | 0.032646 | 1 | 50 | 50 | 46.6028 | 46.6028 | -7.2897
0.1 | Used | 02/07/2026 08:54:23 | 5.0938 | 0.041154 | 1 | 100 | 100 | 94.4841 | 94.4841 | -5.838
0.5 | Used | 02/07/2026 08:58:50 | 26.122 | 0.79071 | 1 | 500 | 500 | 484.8534 | 484.8534 | -3.1239
1 | Used | 02/07/2026 09:03:17 | 53.2945 | 0.62694 | 1 | 1000 | 1000 | 989.2856 | 989.2856 | -1.083
5 | Not used | 02/07/2026 09:07:45 | 286.6444 | 7.568 | 1 | 5000 | 5000
10 | Used | 02/07/2026 09:12:14 | 538.7786 | 0.15297 | 1 | 10000 | 10000 | 10001.8628 | 10001.8628 | 0.018625
Blank | Used | 02/07/2026 08:41:04 | 0.0044152 | 0.00042037 | 100 | 0 | 0 | 0.0033886 | 0.0033886 | 100
Li 670.780
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
Blank | Used | 02/07/2026 08:41:04 | 0.006055 | 0.00023016 | 100 | 0 | 0 | 0.0033948 | 0.0033948 | 100
0.01 | Used | 02/07/2026 08:45:30 | 0.49161 | 0.0077137 | 1 | 10 | 10 | 9.0162 | 9.0162 | -10.9109
0.05 | Used | 02/07/2026 08:49:56 | 2.516 | 0.03256 | 1 | 50 | 50 | 46.5932 | 46.5932 | -7.3117
0.1 | Used | 02/07/2026 08:54:23 | 5.0954 | 0.041335 | 1 | 100 | 100 | 94.4715 | 94.4715 | -5.852
0.5 | Used | 02/07/2026 08:58:50 | 26.1263 | 0.79065 | 1 | 500 | 500 | 484.8411 | 484.8411 | -3.1266
1 | Used | 02/07/2026 09:03:17 | 53.3016 | 0.62643 | 1 | 1000 | 1000 | 989.2641 | 989.2641 | -1.0852
5 | Not used | 02/07/2026 09:07:45 | 286.6813 | 7.5696 | 1 | 5000 | 5000
10 | Used | 02/07/2026 09:12:14 | 538.8476 | 0.15266 | 1 | 10000 | 10000 | 10001.8657 | 10001.8657 | 0.018653
Mn 257.611
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.10664 | 0.0011408 | 1 | 10 | 10 | 10.6484 | 10.6484 | 6.0888
0.05 | Used | 02/07/2026 08:49:56 | 0.53139 | 0.0010032 | 1 | 50 | 50 | 53.5191 | 53.5191 | 6.5753
0.1 | Used | 02/07/2026 08:54:23 | 1.0551 | 0.0024955 | 1 | 100 | 100 | 106.3798 | 106.3798 | 5.9972
0.5 | Used | 02/07/2026 08:58:50 | 5.1467 | 0.055797 | 1 | 500 | 500 | 519.3442 | 519.3442 | 3.7247
1 | Used | 02/07/2026 09:03:17 | 9.9849 | 0.051358 | 1 | 1000 | 1000 | 1007.674 | 1007.674 | 0.76156
5 | Used | 02/07/2026 09:07:45 | 50.3687 | 0.16909 | 1 | 5000 | 5000 | 5083.6554 | 5083.6554 | 1.6456
10 | Used | 02/07/2026 09:12:14 | 98.6368 | 0.24903 | 1 | 10000 | 10000 | 9955.4081 | 9955.4081 | -0.44792
Blank | Used | 02/07/2026 08:41:04 | 0.00083155 | 0.00011172 | 50 | 0 | 0 | -0.030652 | -0.030652 | 100
Mn 259.373
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.076579 | 0.0012554 | 1 | 10 | 10 | 10.3159 | 10.3159 | 3.0619
0.05 | Used | 02/07/2026 08:49:56 | 0.38476 | 0.0029603 | 1 | 50 | 50 | 51.8976 | 51.8976 | 3.6564
0.1 | Used | 02/07/2026 08:54:23 | 0.76815 | 0.0037421 | 1 | 100 | 100 | 103.6268 | 103.6268 | 3.4999
0.5 | Used | 02/07/2026 08:58:50 | 3.7706 | 0.013441 | 1 | 500 | 500 | 508.7315 | 508.7315 | 1.7163
1 | Used | 02/07/2026 09:03:17 | 7.3225 | 0.039455 | 1 | 1000 | 1000 | 987.981 | 987.981 | -1.2165
5 | Used | 02/07/2026 09:07:45 | 37.0676 | 0.2781 | 1 | 5000 | 5000 | 5001.3906 | 5001.3906 | 0.027804
10 | Not used | 02/07/2026 09:12:14 | 72.7932 | 0.19492 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.00011207 | 3.4073e-05 | 50 | 0 | 0 | -0.0015774 | -0.0015774 | 100
Ni 221.648
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.011516 | 0.00048872 | 1 | 10 | 10 | 10.1135 | 10.1135 | 1.1222
0.05 | Used | 02/07/2026 08:49:56 | 0.056965 | 0.0019207 | 1 | 50 | 50 | 50.9917 | 50.9917 | 1.9449
0.1 | Used | 02/07/2026 08:54:23 | 0.11614 | 0.0020647 | 1 | 100 | 100 | 104.2121 | 104.2121 | 4.0419
0.5 | Used | 02/07/2026 08:58:50 | 0.56622 | 0.0017491 | 1 | 500 | 500 | 509.0337 | 509.0337 | 1.7747
1 | Used | 02/07/2026 09:03:17 | 1.1064 | 0.0037153 | 1 | 1000 | 1000 | 994.8846 | 994.8846 | -0.51417
5 | Not used | 02/07/2026 09:07:45 | 5.5953 | 0.041439 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 11.0727 | 0.072064 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.0002678 | 4.6747e-05 | 50 | 0 | 0 | -0.0036943 | -0.0036943 | 100
Ni 231.604
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.0068746 | 0.00039145 | 1 | 10 | 10 | 10.1995 | 10.1995 | 1.9562
0.05 | Used | 02/07/2026 08:49:56 | 0.033828 | 0.0010586 | 1 | 50 | 50 | 51.7372 | 51.7372 | 3.3577
0.1 | Used | 02/07/2026 08:54:23 | 0.068152 | 0.0011775 | 1 | 100 | 100 | 104.6324 | 104.6324 | 4.4273
0.5 | Used | 02/07/2026 08:58:50 | 0.33166 | 0.00036531 | 1 | 500 | 500 | 510.7217 | 510.7217 | 2.0993
1 | Used | 02/07/2026 09:03:17 | 0.6486 | 0.0015369 | 1 | 1000 | 1000 | 999.1517 | 999.1517 | -0.084905
5 | Used | 02/07/2026 09:07:45 | 3.2594 | 0.016987 | 1 | 5000 | 5000 | 5022.5465 | 5022.5465 | 0.4489
10 | Used | 02/07/2026 09:12:14 | 6.4815 | 0.042026 | 1 | 10000 | 10000 | 9988.1414 | 9988.1414 | -0.11873
Blank | Used | 02/07/2026 08:41:04 | 0.00024913 | 8.9837e-05 | 50 | 0 | 0 | -0.010852 | -0.010852 | 100
Ni 232.003
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.0022767 | 5.9729e-05 | 1 | 10 | 10 | 10.1108 | 10.1108 | 1.0957
0.05 | Used | 02/07/2026 08:49:56 | 0.012252 | 0.00023467 | 1 | 50 | 50 | 51.318 | 51.318 | 2.5683
0.1 | Used | 02/07/2026 08:54:23 | 0.024585 | 0.00040723 | 1 | 100 | 100 | 102.2655 | 102.2655 | 2.2153
0.5 | Used | 02/07/2026 08:58:50 | 0.12226 | 0.0011433 | 1 | 500 | 500 | 505.768 | 505.768 | 1.1404
1 | Used | 02/07/2026 09:03:17 | 0.23711 | 0.0016953 | 1 | 1000 | 1000 | 980.2285 | 980.2285 | -2.017
5 | Used | 02/07/2026 09:07:45 | 1.2075 | 0.0022266 | 1 | 5000 | 5000 | 4988.7935 | 4988.7935 | -0.22463
10 | Used | 02/07/2026 09:12:14 | 2.4223 | 0.0087416 | 1 | 10000 | 10000 | 10007.2018 | 10007.2018 | 0.071966
Blank | Used | 02/07/2026 08:41:04 | -0.00016949 | 3.4914e-05 | 50 | 0 | 0 | 0.0057256 | 0.0057256 | 100
Pb 220.353
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.0014558 | 6.4449e-05 | 1 | 10 | 10 | 9.4674 | 9.4674 | -5.6255
0.05 | Used | 02/07/2026 08:49:56 | 0.0091912 | 0.00038193 | 1 | 50 | 50 | 50.4711 | 50.4711 | 0.93339
0.1 | Used | 02/07/2026 08:54:23 | 0.018944 | 0.00070303 | 1 | 100 | 100 | 102.1679 | 102.1679 | 2.1219
0.5 | Not used | 02/07/2026 08:58:50 | 0.096549 | 0.00057961 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 0.18828 | 0.0015449 | 1 | 1000 | 1000 | 999.7597 | 999.7597 | -0.024034
5 | Not used | 02/07/2026 09:07:45 | 0.96128 | 0.0076705 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 1.9085 | 0.016593 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | -0.00033042 | 4.5082e-05 | 50 | 0 | 0 | -0.00074645 | -0.00074645 | 100
Pb 283.305
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.0010536 | 0.00023541 | 1 | 10 | 10
0.05 | Not used | 02/07/2026 08:49:56 | 0.0047329 | 4.1928e-06 | 1 | 50 | 50
0.1 | Not used | 02/07/2026 08:54:23 | 0.0096792 | 0.00029365 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 0.048628 | 0.0002856 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 0.095032 | 0.00096712 | 1 | 1000 | 1000 | 983.6465 | 983.6465 | -1.6625
5 | Used | 02/07/2026 09:07:45 | 0.4994 | 0.0055022 | 1 | 5000 | 5000 | 4906.9158 | 4906.9158 | -1.897
10 | Used | 02/07/2026 09:12:14 | 1.0291 | 0.012096 | 1 | 10000 | 10000 | 10046.6725 | 10046.6725 | 0.46456
Blank | Used | 02/07/2026 08:41:04 | 0.00011833 | 0.00013647 | 1 | 0 | 0 | 62.7653 | 62.7653 | 100
Se 196.090
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.00017704 | 0.00013907 | 1 | 10 | 10 | 13.1496 | 13.1496 | 23.952
0.05 | Used | 02/07/2026 08:49:56 | 0.0032859 | 0.00015805 | 1 | 50 | 50 | 56.7743 | 56.7743 | 11.932
0.1 | Used | 02/07/2026 08:54:23 | 0.0069098 | 0.00036936 | 1 | 100 | 100 | 107.6254 | 107.6254 | 7.0851
0.5 | Not used | 02/07/2026 08:58:50 | 0.034716 | 6.8037e-05 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 0.06847 | 0.00013341 | 1 | 1000 | 1000 | 971.4511 | 971.4511 | -2.9388
5 | Used | 02/07/2026 09:07:45 | 0.35594 | 0.00085366 | 1 | 5000 | 5000 | 5005.2853 | 5005.2853 | 0.10559
10 | Not used | 02/07/2026 09:12:14 | 0.70841 | 0.0029837 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | -0.00035283 | 8.5564e-05 | 1 | 0 | 0 | 5.7142 | 5.7142 | 100
Se 204.050
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | -0.00085437 | 7.059e-05 | 1 | 10 | 10
0.05 | Not used | 02/07/2026 08:49:56 | 0.00086816 | 0.00013205 | 1 | 50 | 50
0.1 | Not used | 02/07/2026 08:54:23 | 0.0027595 | 6.2793e-05 | 1 | 100 | 100
0.5 | Used | 02/07/2026 08:58:50 | 0.018945 | 9.4415e-05 | 1 | 500 | 500 | 488.8392 | 488.8392 | -2.2831
1 | Used | 02/07/2026 09:03:17 | 0.03828 | 0.00014627 | 1 | 1000 | 1000 | 963.0442 | 963.0442 | -3.8374
5 | Used | 02/07/2026 09:07:45 | 0.2037 | 0.0012754 | 1 | 5000 | 5000 | 5020.1048 | 5020.1048 | 0.40049
10 | Used | 02/07/2026 09:12:14 | 0.4065 | 0.00060781 | 1 | 10000 | 10000 | 9994.0081 | 9994.0081 | -0.059955
Blank | Used | 02/07/2026 08:41:04 | -0.00098568 | 6.0651e-05 | 50 | 0 | 0 | 0.013601 | 0.013601 | 100
Sr 407.771
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 1.7329 | 0.019816 | 1 | 10 | 10 | 10.4708 | 10.4708 | 4.4959
0.05 | Used | 02/07/2026 08:49:56 | 8.8845 | 0.097897 | 1 | 50 | 50 | 53.8026 | 53.8026 | 7.0677
0.1 | Used | 02/07/2026 08:54:23 | 17.6646 | 0.069577 | 1 | 100 | 100 | 107.0017 | 107.0017 | 6.5436
0.5 | Not used | 02/07/2026 08:58:50 | 84.7699 | 0.54215 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 164.8894 | 3.3266 | 1 | 1000 | 1000 | 999.0403 | 999.0403 | -0.09606
5 | Not used | 02/07/2026 09:07:45 | 672.4405 | 6.4425 | 1 | 5000 | 5000
10 | Not used | 02/07/2026 09:12:14 | 670.9793 | 5.4231 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.0040487 | 0.002007 | 50 | 0 | 0 | -0.0041262 | -0.0041262 | 100
Sr 421.552
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Not used | 02/07/2026 08:45:30 | 0.63122 | 0.0083311 | 1 | 10 | 10
0.05 | Not used | 02/07/2026 08:49:56 | 3.1523 | 0.029225 | 1 | 50 | 50
0.1 | Not used | 02/07/2026 08:54:23 | 6.2216 | 0.034358 | 1 | 100 | 100
0.5 | Not used | 02/07/2026 08:58:50 | 29.8083 | 0.7047 | 1 | 500 | 500
1 | Used | 02/07/2026 09:03:17 | 58.1069 | 1.0802 | 1 | 1000 | 1000 | 1002.3958 | 1002.3958 | 0.23901
5 | Used | 02/07/2026 09:07:45 | 283.1794 | 5.3598 | 1 | 5000 | 5000 | 5123.1195 | 5123.1195 | 2.4032
10 | Used | 02/07/2026 09:12:14 | 546.0517 | 3.4896 | 1 | 10000 | 10000 | 9935.8962 | 9935.8962 | -0.64517
Blank | Used | 02/07/2026 08:41:04 | 0.0021029 | 0.0011308 | 1 | 0 | 0 | -61.4115 | -61.4115 | 100
Zn 213.856
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.024102 | 0.00064907 | 1 | 10 | 10 | 10.0481 | 10.0481 | 0.47822
0.05 | Used | 02/07/2026 08:49:56 | 0.11312 | 0.0010075 | 1 | 50 | 50 | 49.9717 | 49.9717 | -0.056624
0.1 | Used | 02/07/2026 08:54:23 | 0.22693 | 0.002795 | 1 | 100 | 100 | 101.016 | 101.016 | 1.0058
0.5 | Used | 02/07/2026 08:58:50 | 1.107 | 0.0030405 | 1 | 500 | 500 | 495.7268 | 495.7268 | -0.86201
1 | Not used | 02/07/2026 09:03:17 | 2.17 | 0.00575 | 1 | 1000 | 1000
5 | Used | 02/07/2026 09:07:45 | 11.1508 | 0.067838 | 1 | 5000 | 5000 | 5000.4033 | 5000.4033 | 0.0080652
10 | Not used | 02/07/2026 09:12:14 | 22.3436 | 0.094024 | 1 | 10000 | 10000
Blank | Used | 02/07/2026 08:41:04 | 0.0017011 | 7.3095e-05 | 50 | 0 | 0 | 0.0011337 | 0.0011337 | 100
Zn 206.200
Standard | Use | Measurement Date Time | Intensity | Intensity SD | Weight | Cert. Conc [ppb] | Prep. Corr. Cert. Conc. [ppb] | Calc. Conc. [ppb] | Calc. Conc. Uncorr. [ppb] | Rel. Conc. Deviation[%]
0.01 | Used | 02/07/2026 08:45:30 | 0.015197 | 0.00074734 | 1 | 10 | 10 | 8.6814 | 8.6814 | -15.1889
0.05 | Used | 02/07/2026 08:49:56 | 0.073291 | 0.0027455 | 1 | 50 | 50 | 48.7306 | 48.7306 | -2.6048
0.1 | Used | 02/07/2026 08:54:23 | 0.14957 | 0.0033399 | 1 | 100 | 100 | 101.3178 | 101.3178 | 1.3006
0.5 | Used | 02/07/2026 08:58:50 | 0.73212 | 0.0045265 | 1 | 500 | 500 | 502.9161 | 502.9161 | 0.57984
1 | Not used | 02/07/2026 09:03:17 | 1.4315 | 0.014143 | 1 | 1000 | 1000
5 | Not used | 02/07/2026 09:07:45 | 7.3534 | 0.10178 | 1 | 5000 | 5000
10 | Used | 02/07/2026 09:12:14 | 14.5081 | 0.17311 | 1 | 10000 | 10000 | 9999.8471 | 9999.8471 | -0.0015291
Blank | Used | 02/07/2026 08:41:04 | 0.00043843 | 5.4255e-05 | 1 | 0 | 0 | -1.493 | -1.493 | 100

## Instrument State
Sample Name | CCV 1
Measurement Date | 02/07/2026 13:26:07
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 34.2 °C (Max 58.0)
Power Amplifier Temperature | 60.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | RE_AC_452
Measurement Date | 02/07/2026 10:01:21
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.3 °C (Max 58.0)
Power Amplifier Temperature | 59.8 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Blank
Measurement Date | 02/07/2026 09:30:09
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 58.7 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV 1
Measurement Date | 02/07/2026 09:34:35
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.3 °C (Max 58.0)
Power Amplifier Temperature | 58.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV_50
Measurement Date | 02/07/2026 09:39:03
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | MB
Measurement Date | 02/07/2026 09:43:32
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | STOCK_10
Measurement Date | 02/07/2026 09:47:59
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.5 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | LFS_1/10
Measurement Date | 02/07/2026 09:52:27
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.5 °C (Max 58.0)
Power Amplifier Temperature | 58.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | IHS_54_1/50
Measurement Date | 02/07/2026 09:56:54
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2000 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | 55429_1554316
Measurement Date | 02/07/2026 10:05:48
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 59.8 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_as is
Measurement Date | 02/07/2026 10:10:15
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.6 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_dup
Measurement Date | 02/07/2026 10:14:42
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.6 °C (Max 58.0)
Power Amplifier Temperature | 58.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_ms_Na
Measurement Date | 02/07/2026 10:19:10
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_ms_K
Measurement Date | 02/07/2026 10:23:38
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.4 °C (Max 58.0)
Power Amplifier Temperature | 59.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_ms_Mg+Ca
Measurement Date | 02/07/2026 10:28:06
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.5 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_ms_met_100 ppb
Measurement Date | 02/07/2026 10:32:34
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 31.9 °C (31.0 - 33.0)
Torch Box Temperature | 33.5 °C (Max 58.0)
Power Amplifier Temperature | 58.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_ms_met_100 ppb_d
Measurement Date | 02/07/2026 10:37:03
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.6 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Var_9_ms_met_200 ppb_B
Measurement Date | 02/07/2026 10:41:32
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 59.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Blank
Measurement Date | 02/07/2026 10:46:00
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV 1
Measurement Date | 02/07/2026 10:50:26
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.6 °C (Max 58.0)
Power Amplifier Temperature | 58.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV_50
Measurement Date | 02/07/2026 10:54:54
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 31.9 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 60.1 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | LFS_1/100
Measurement Date | 02/07/2026 10:59:24
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.6 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2001 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Sis_Al+Fe
Measurement Date | 02/07/2026 11:03:51
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 59.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_459_1/10
Measurement Date | 02/07/2026 11:08:16
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.6 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | MA_54
Measurement Date | 02/07/2026 11:12:43
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 60.1 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | MA_54_Dup
Measurement Date | 02/07/2026 11:17:09
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 59.5 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | MA_54_ms_Na+k
Measurement Date | 02/07/2026 11:21:36
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | MA_54_ms_Mg+Ca
Measurement Date | 02/07/2026 11:26:03
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 59.6 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | MA_54_ms_B_300 ppb
Measurement Date | 02/07/2026 11:30:30
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 59.8 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Sis_Mg
Measurement Date | 02/07/2026 11:34:58
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 59.3 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Sis_Ca
Measurement Date | 02/07/2026 11:39:25
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | blank
Measurement Date | 02/07/2026 11:43:53
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 31.9 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 60.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | blank
Measurement Date | 02/07/2026 11:48:18
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 60.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | Blank
Measurement Date | 02/07/2026 11:52:46
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 59.3 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV 1
Measurement Date | 02/07/2026 11:57:12
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 59.3 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2002 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV_50
Measurement Date | 02/07/2026 12:01:39
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 60.5 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | LFS_1/100
Measurement Date | 02/07/2026 12:06:07
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.8 °C (Max 58.0)
Power Amplifier Temperature | 59.5 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | IHS_53_1/5
Measurement Date | 02/07/2026 12:10:35
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | IHS_54_1/2
Measurement Date | 02/07/2026 12:15:04
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.7 °C (Max 58.0)
Power Amplifier Temperature | 60.1 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | IHS_55_1/5
Measurement Date | 02/07/2026 12:19:32
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 60.1 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_455_1/5
Measurement Date | 02/07/2026 12:23:58
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_456_1/5
Measurement Date | 02/07/2026 12:28:24
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_457_1/5
Measurement Date | 02/07/2026 12:32:50
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 34.0 °C (Max 58.0)
Power Amplifier Temperature | 60.1 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_458_1/5
Measurement Date | 02/07/2026 12:37:17
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 31.9 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_459_1/2
Measurement Date | 02/07/2026 12:41:43
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 34.0 °C (Max 58.0)
Power Amplifier Temperature | 59.5 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | AC_460_1/2
Measurement Date | 02/07/2026 12:46:11
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 12:50:38
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 12:55:04
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 12:59:30
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 62.0 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2003 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 13:03:56
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.9 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 13:08:23
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 60.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 13:12:49
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 33.9 °C (Max 58.0)
Power Amplifier Temperature | 59.6 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 13:17:15
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 34.1 °C (Max 58.0)
Power Amplifier Temperature | 60.3 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | BLANK
Measurement Date | 02/07/2026 13:21:40
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 34.2 °C (Max 58.0)
Power Amplifier Temperature | 60.3 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13
Sample Name | CCV 1
Measurement Date | 02/07/2026 13:26:07
Nebulizer Pressure | 2.6 bar (2.0 - 4.0)
Main Argon Pressure | 7.0 bar (6.0 - 8.0)
Light Tube Flow | 0.90 l/min
Twin Interface Radial Flow | 0.70 l/min
Dark Current Measurement State | Valid
UV Plus Purifier UTC Date and Time | יום ראשון 01 דצמבר 2024 (-577 days)
Optics Temperature | 32.0 °C (31.0 - 33.0)
Torch Box Temperature | 34.2 °C (Max 58.0)
Power Amplifier Temperature | 60.2 °C (Max 100.0)
Door State | Closed
Exhaust State | Working
Torch Box Cover | Closed
Generator Error | No error
Waste Container | Ok
Instrument Number | 157257
Instrument Type | FMT46
Generator OperationHours | 2004 hours
Spectra Calibration Date | 01/07/2026 09:42:13