05-SOP

INDUCTIVELY COUPLED PLASMA-OPTICAL EMISSION SPECTROMETRY

BASED ON EPA METHOD 6010D

|  | SCOPE AND APPLICATION | # 2 |
| --- | --- | --- |
|  | # SUMMARY OF METHOD | # 2 |
|  | # DEFINITIONS | # 3 |
|  | INTERFERENCES | # 3 |
|  | # SAFETY | # 4 |
|  | # EQUIPMENT AND SUPPLIES | # 4 |
|  | # REAGENTS AND STANDARDS | # 5 |
|  | # SAMPLE COLLECTION, PRESERVATION, AND HANDLING | # 6 |
|  | # QUALITY CONTROL | # 6 |
|  | # CALIBRATION AND STANDARDIZATION | # 7 |
|  | # PROCEDURE | # 9 |
|  | # DATA ANALYSIS AND CALCULATIONS | # 9 |
|  | # WASTE MANAGEMENT | # 9 |
|  | # REFERENCES | # 9 |
|  | # TABLES, APPENDICES | # 11 |
|  | # Appendix 1. PREPARATION PROCEDURE OF STANDARDS AND QC SAMPLES | # 12 |
|  | # Appendix 2. EXAMPLE FOR SEQUENCE FORMATION | # 17 |

SCOPE AND APPLICATION

  - Inductively coupled plasma – optical emission spectrometry (ICP-OES) is a technique used to determine trace elements in aqueous solutions. The method is applicable to the direct analysis of the dissolved constituents (elements are listed in the table below) in relatively clean aqueous matrices, e.g. drinking, surface, and ground water. Water samples should be filtered and acidified before analysis. Other elements and water matrices may be analyzed by this method if acceptable performance at the concentration of interest is demonstrated.

| Element | Symbol | CASRN |
| --- | --- | --- |
| Arsenic | As | 7440-38-2 |
| Boron | B | 7440-42-8 |
| Calcium | Ca | 7440-70-2 |
| Chromium | Cr | 7440-47-3 |
| Copper Iron Lead | Cu Fe Pb | 7440-50-8 7439-89-6 7439-92-1 |
| Magnesium | Mg | 7439-95-4 |
| Manganese | Mn | 7439-96-5 |
| Potassium | K | 7440-09-7 |
| Sodium | Na | 7440-23-5 |

  - Use of this method is restricted to spectroscopists who are knowledgeable in the correction of spectral, chemical, and physical interferences described in this method. Analyst should understand the data quality objectives (DQO) prior to analysis.

SUMMARY OF METHOD

  - This method describes multi-element determinations by ICP-OES using simultaneous optical system and axial and radial viewing of the plasma. The instrument measures characteristic emission spectra by optical spectrometry. Samples are nebulized and the resulting aerosol is transported to the plasma torch. Element-specific emission spectra are produced by a radio frequency inductively coupled plasma. The spectra are dispersed by a grating spectrometer, and the intensities of the emission lines are monitored by photosensitive devices.

  - The LLOQ of the elements are: Na 1 mg/L; K 0.5 mg/L; Mg 0.1 mg/L; Ca 0.5 mg/L; As 0.05 mg/L; B 0.05 mg/L; Cr 0.05 mg/L; Cu 0.05 mg/L; Fe 0.05 mg/L; Mn 0.05 mg/L; Pb 0.05 mg/L.

DEFINITIONS

| ICP-OES | Inductively Coupled Plasma-Optical Emission Spectrometry |
| --- | --- |
| SDS | Safety Data Sheet |
| ICV | Initial Calibration Verification |
| CCV | Continuing Calibration Verification |
| ICB | Initial Calibration Blank |
| CCB | Continuing Calibration Blank |
| DQO | Data Quality Objectives |
| QC | Quality Control |
| MS/MSD | Matrix spike/matrix spike duplicate |
| IDL | Instrument Detection Limit |
| LLOQ SIC | Lower Limit of Quantitation Spectral interference check |

INTERFERENCES

  - Spectral interferences can be caused by background emission from continuous or recombination phenomena and/or stray light from the line emission of high concentration elements, spectral line overlaps between target elements, unresolved overlap of molecular. Spectral overlaps may be avoided by using an alternate wavelength for quantitation and by measuring SIC solution (see Sec ‎9.9). More extensive information on interfering effects at various wavelengths and resolutions is available in instrument software.

  - Physical interferences are effects associated with the sample nebulization and transport processes. Changes in viscosity and surface tension can cause significant inaccuracies, especially in samples containing high dissolved solids or acid concentrations. If physical interferences are present, they must be reduced by diluting the sample, by using a peristaltic pump, by using an internal standard, or by using a high solids nebulizer.

  - Chemical interferences include molecular compound formation, ionization effects, and solute vaporization effects. This can be minimized by careful selection of operating conditions, buffering the sample, by standard additions method, and using internal standard (1 mg/L Yttrium standard solution is used for this purpose, refer to Sec ‎7.5.5 and Table 2 in Appendix 1(. The internal standard element intensity is used by the instrument to ratio the analyte intensity signals for both calibration and quantitation.

  - Memory interferences result when analytes in a previous sample contribute to the signals measured in a new sample. The site where these effects occur is dependent on the element (see Sec ‎12.2) and can be minimized by flushing the system with a rinse blank between samples. A rinse period of at least 60 seconds should be used between the samples and standards.

  - Reagents and sample processing may yield artifacts and/or interferences to sample analysis. All of these materials must be demonstrated to be free from interferences under the conditions of the analysis by analyzing method blanks.

SAFETY

  - This method does not address all safety issues associated with its use. The laboratory is responsible for maintaining a safe work environment and a current awareness file of OSHA regulations regarding the safe handling of the chemicals specified in this method. A reference file of safety data sheets (SDS) should be available to all personnel involved in these analyses. For more information, refer to "Laboratory Safety Working Procedure" W-03.

  - Concentrated nitric acid is moderately toxic and extremely irritating to skin and mucus membranes. Use these reagents in a hood and if eye or skin contact occurs, flush with large volumes of water.

EQUIPMENT AND SUPPLIES

  - Inductively coupled, argon plasma, emission spectrometer.

    - Computer-controlled, emission spectrometer with background correction capability.

    - Radio-frequency generator compliant with FCC regulations.

    - Mass flow controller for argon nebulizer gas supply.

    - Peristaltic pump.

    - Auto sampler.

    - Argon gas supply – high purity.

  - Volumetric flasks of suitable precision and accuracy.

  - Volumetric pipettes of suitable precision and accuracy.

  - 10 ml Plastic disposable syringes (Medi Plus, P/N 7290010993659 or equivalent).

  - 0.22 micron PVDF syringe filter (Millex, P/N SLGV033RS or equivalent).

  - PP Centrifuge Tubes with Plug Seal Cap, Sterile 15 ml (Corning, P/N WV-430790 or equivalent)

REAGENTS AND STANDARDS

  - Reagent-grade or ultra-high purity grade chemicals are used in all tests. Other grades may be used, provided it is first ascertained that the reagent is of sufficiently high purity to permit its use without lessening the accuracy of the determination.

  - Nitric acid (65%), Suprapur (Merck 1.00441.1000 or equivalent)

  - Nitric acid (65%) for rinse Blank, Emsure (Merck 1.01799.1000 or equivalent)

  - De-ionized water (DIW), Type I reagent grade, with a resistance of 18 MΩ-cm or better.

  - Standard stock solutions:

    - Calcium standard stock solution, 1000 mg/L (Merck 1.19778.0500 or equivalent).

    - Sodium standard stock solution, 1000 mg/L (Merck 1.19507.0500 or equivalent).

    - Magnesium standard stock solution, 1000 mg/L (Merck 1.19788.0500 or equivalent).

    - Potassium standard stock solution, 1000 mg/L (Merck 1.70230.0500 or equivalent).

    - Yttrium standard stock solution, 1000 mg/L (Merck 1.70368.0100 or equivalent).

    - ICP multi-element standard solution VIII, 100 mg/L (Merck 1.09492.0100 or equivalent).

    - ICV for cations: ICP multi-element standard solution 10 (Merck 1.09493.0100 or equivalent).

    - ICV for metals: Multi-element standard WW LFS solution #1 (Inorganic Ventures, part number WW-LFS-1-125mL or equivalent).

    - ICP Wavelength Calibration Solution, I CAL :  10 mg/L Ca,Cu,Eu,Fe,In,K,Ni,P,Si,Ti,V,Y,Zr, 5 mg/LMn and 2 mg/LMn Be,Li and Sr.  (OrganicVenture, PN genesis-ical-125ML or equivalent).

For the preparation procedure of calibration and QC solutions refer to Sec. 10.0.

SAMPLE COLLECTION, PRESERVATION, AND HANDLING

  - Sample should be filtered and acidified up to 72 hours from the date of sample arrival to the laboratory (see Sec. 11.1).  For more instructions, see SW 846, Chapter Three, Inorganic Analytes, for sample collection and preservation instructions.

QUALITY CONTROL

  - Instrument detection limit (IDL) is useful mean to evaluate the instrument noise level and response changes over time for each analyte. IDLs should be determined at least once using new equipment, after major instrument maintenance, and/or at a frequency designated by the project.

  - Initial Calibration blank (ICB) is used to establish the analytical curve and to verify the validity of the calibration curve during the sequence (CCB). Method blank (MB) is run immediately after the calibration curve and is used to identify possible contamination resulting from either the reagents (acid) or the equipment (filters) used during sample processing. For each batch of samples analyzed, method blank must be carried throughout the entire sample preparation and instrument determination process. Rinse blank is used to rinse the system between samples. Blanks are considered to be acceptable if target analyte concentrations are less than 1/2 of the lower limit of quantitation (LLOQ).

  - The linear range (LDR) establishes the highest concentration that may be reported without diluting the sample. For Ca, Na, Mg and K the linear range is up to 500 ppm and must be verify each 6 months. LDR are considered to be acceptable if the recovery is ± 10%.

  - Matrix spike/matrix spike duplicates (MS/MSDs) pair is used to document the bias and precision of a method in a given water sample matrix.

    - For each new project (and once a year for an existing project), (i) one MS and one duplicate unspiked sample or (ii) one matrix spike/matrix spike duplicate (MS/MSD) pair should be included for documenting the effect of the matrix (e.g. ground, surface, drinking water). If samples are expected to contain target analytes, then MS and a duplicate analysis of an unspiked field sample should be analyzed. If samples are not expected to contain target analytes, the laboratory should use an MS/MSD pair. MS/MSD samples should be spiked with each target element between the low- and mid-level standards. Acceptance criteria is set at ± 25% recovery of the target element spiked value for accuracy and ≤20 relative percent difference (RPD) for precision.

    - The RPD between duplicate or MS determinations is calculated as follows:

RPD=D1-D2(D1+D2)/2*100,

where:

RPD = relative percent difference,

D1 = MS or first sample analysis value,

D2 = MSD or duplicate sample analysis value.

  - Lower Limit of Quantitation (LLOQ) check standard. The laboratory should establish the LLOQ as the lowest point of quantitation, which is the lowest concentration in the calibration curve. The LLOQ is initially verified by the analysis of at least 7 replicate samples, spiked at the LLOQ and processed through all preparation and analysis steps of the method. The mean recovery and relative standard deviation of these samples provide an initial statement of precision and accuracy at the LLOQ. The mean recovery should be ±20% of the true value and RSD should be <10%. LLOQ should be verified quarterly.

  - Initial Calibration Verification (ICV). The ICV standard is prepared from a separate source than the initial calibration standards curve (see Sec‎10.5, and Table 2). It is analyzed at approximately the mid-level of the calibration and serves as a check that the initial calibration standards are at the correct concentrations. The acceptance range is ±10% of the true value.

  - Continues calibration verification (CCV). standard should be prepared using the same acid matrix and stock standards employed when preparing the calibration standards, CCV should contain concentrations equal or nearly equivalent to the midpoint concentration of the calibration curve (see Sec 10.6 and Table 2).

  - The calibration curve must be verified at the end of each analysis batch run and after every 10 samples through the analysis of CCV standard and a CCB. For the curve to be considered valid, for the analysis result of the CCV standard must be within ±10% of its true value and the CCB must not contain target analytes above the LLOQ. If the calibration cannot be verified, sample analysis must be discontinued, the cause of the problem determined and the instrument recalibrated. All samples following the last acceptable CCV standard must be reanalyzed.

  - SIC solutions must verify the absence of spectral interference. The SIC solution should be prepared using reagent water fortified by: Al 10 mg/L; Mg 1000 mg/L; Fe 10 mg/L; Ca 1000 mg/L. (see Sec 10.7 and Table 2)

SIC solutions must run quarterly. Concentration considered to be acceptable if all analytes concentrations are less than 1/2 of the LLOQ of each element.

  - IS recovery is calculated automatically by the software and shown at "Y 224.3 nm" label in results tab. It is considered to be acceptable if the recovery is between 70-130%. If it is not acceptable, reanalyze the solution or check the calibration curve.

CALIBRATION AND STANDARDIZATION

  - Reagent water is used for all standard preparations and dilution to the final volume in volumetric flasks. High purity nitric acid (HNO3, 65%, Suprapure or equivalent) is used to acidify all calibration standards, blanks, ICV at a level of 5% v/v.

  - Internal standard preparation

    - Stock solution of internal Standard (IS), 20 mg/L, is prepared by aliquoting 2 ml of Yttrium standard stock solution (1000 mg/L) into 100 ml volumetric flask.

    - To achieve 1ppm IS in the solution, add 500 µl to 10ml sample or 2.5 ml to 50ml standard from stock solution of 20 mg/L IS.

    - Internal standard 1 mg/L is added prior to analysis on the day of preparation to all standards and samples (nitric acid and IS is added above the final volume of the solution in the corning tube).

  - Calibration Standard Solutions - Low concentrations:

All calibration standards (low and high concentrations) are prepared into 50 ml volumetric flask. At the end of the standards preparation, transfer into Corning 50 ml tube, add 2.5 mL of nitric acid (SupraPure 65%) and 2.5 mL from 20ppm IS solution for each standard.

    - Calibration Standard Stock Solution 10 mg/L is prepared by aliquoting 5 ml from the ICP multi-element standard VIII standard and 500µl of Arsenic standard stock solution (1000 mg/L) into 50 ml volumetric flask. Valid for one month.

    - Calibration Standard 5 mg/L is prepared by aliquoting 2.5 ml from the ICP multi-element standard VIII standard and 250µl of Arsenic standard stock solution (1000 mg/L) into 50 ml volumetric flask. Valid for one month.

    - Calibration Standard 1 mg/L is prepared by aliquoting 1 ml from the ICP multi-element standard VIII standard and 100µl of Arsenic standard stock solution (1000 mg/L) into 100 ml volumetric flask. Valid for one month.

    - Calibration Standard 0.5 mg/L is prepared by aliquoting 250µl from the ICP multi-element standard VIII standard and 25µl of Arsenic standard stock solution (1000 mg/L) into 50 ml volumetric flask. Valid for one month.

    - Calibration Standard 0.1 mg/L (with 0.5 mg/L K) is prepared by aliquoting 5 ml of 1 mg/L standard solution into 50 ml volumetric flask. Valid for one month.

    - Calibration Standard 0.05 mg/L (with 0.5 mg/L K) is prepared by aliquoting 2.5 ml of 1 mg/L standard solution into 50 ml volumetric flask. Valid for one month.

    - Calibration Standard 0.01 mg/L (with 0.5 mg/L K) is prepared by aliquoting 500µl of 1 mg/L standard solution into 50 ml volumetric flask. Valid for one month.

  - Calibration Standard Solutions- High concentrations (major elements):

    - Calibration Standard 50 mg/L is prepared by aliquoting 5 ml of the

Calcium, Sodium, Magnesium and Potassium standards stock solutions (1000 mg/L) into 100 ml volumetric flask. Valid for one month.

    - Calibration Standard 100 mg/L is prepared by aliquoting 5 ml of the Calcium, Sodium, Magnesium and Potassium standards stock solutions (1000 mg/L) into 50 ml volumetric flask. Valid for one month.

    - Calibration Standard 200 mg/L is prepared by aliquoting 10 ml from the Calcium, Sodium, Magnesium and Potassium standards stock solutions (1000 mg/L) into 50 ml volumetric flask. Valid for one month.

  - Initial calibration verification (ICV)

    - ICV for Major cations: ICP multi-element standard solution stock-10 (or equivalent). Used without dilution. Pour the solution into 10ml tube, add 500µL of nitric acid (SupraPure 65%) and 500µL from 20ppm IS solution.

    - ICV for Major cations: Multi-element standard WW LFS solution #1 (or equivalent). LFS 1/10 is prepared by aliquoting 5ml of LFS solution #1 standard solution into 50 ml volumetric flask. Transfer into Corning 50 ml tube. add 2.5 mL of nitric acid (SupraPure 65%) and 2.5 mL from 20ppm IS solution for each standard. Valid for one month.

    - ICV for metals: Multi-element standard WW LFS solution #1 (or equivalent). LFS 1/100 is prepared by aliquoting 500µl of LFS solution #1 standard solution into 50 ml volumetric flask. Transfer into Corning 50 ml tube. add 2.5 mL of nitric acid (SupraPure 65%) and 2.5 mL from 20ppm IS solution for each standard. Valid for one month.

  - Continuing calibration verification (CCV):

    - Traces continuing calibration verification (CCV) 1 mg/L. For preparation procedure see Sec. 10.3.3 .

    - Majorscontinuing calibration verification 50 mg/L. For preparation  procedure see Sec. 10.4.1 .

  - Spectra interference solutions (SIC)

    - SIC solution of 10 ppm Al and Fe is prepared by aliquoting 100 µl of Aluminum stock solution (1000 mg/L) and 100 µl of Iron stock solution (1000 mg/L) into 10 ml volumetric flask. Transfer into Corning 10 ml tube. Add 500µl Nitric acid (SupraPure 65%) and 500 µl from 20ppm IS solution.

    - SIC solution of 1000 ppm Mg is prepared by transfer 10 ml of magnesium  stock solution (1000 mg/L) into Corning 10 ml tube. Add 500µl  Nitric acid (SupraPure 65%) and 500 µl from 20ppm IS solution.

  - Blanks

    - Calibration blank is prepared by aliquoting 50 ml reagent water and adding 2.5 ml Nitric acid (SupraPure, 65%) and 2.5ml IS (20 ppm Y stock), to achieve the same concentrations found in the standards and

samples (5% v/v Nitric acid and 1 ppm IS).

    - The method blank is prepared by filtering 10 ml of reagent water and adding 500 µl Nitric acid (SupraPure, 65%) and 500 µl of IS (20ppm Y stock), process equal to that used for actual aqueous samples (5% v/v Nitric acid and 1 ppm IS).

    - Rinse Blank is prepared by acidifying 1 liter of reagent water with 50ml Nitric acid (Emsure 65%, 5% v/v).

10.9 Calibration curve: Table 2 presents all calibration standards, ICV, CCV and other standard solutions, the example of preparation from the corresponding stock solutions and validity period for each standard.

    - A calibration curve must be analyzed daily. The instrument is calibrated using a multipoint calibration curve. A minimum of three standards (including blank) are required and the coefficient of determination (r2) should be >0.990. The low concentration standard must be at or below the LLOQ. Second-order calibration curves may be used for alkali or alkaline earth metal.  the second-order curve fit should have a coefficient of determination (r2) should be >0.990 or better.

    - Calibration standards should be prepared on an as-needed basis. Calibration standard solutions are valid for one month. All standard should be stored in the refrigerator when not in use . If the ICV fails to meet the acceptance criteria, troubleshoot the situation, and then prepare a new set of calibration standards if needed and recalibrate the instrument.

    - After the calibration is completed it is verified using CCB, CCV and ICV.

PROCEDURE

  - Preliminary treatment of water samples includes filtering and acidification with nitric acid at the same level as calibration standards (5% v/v of acid). Groundwater samples which have been prefiltered and acidified will not need acid digestion. Only before the running analysis, 500µl of internal standard (from stock solution of 20 ppm Y) must be added to each 10 ml sample before samples analysis, to achieve 1ppm Y in sample.

  - Set up the instrument using appropriate operating conditions established as detailed in the W-12 (sec. 1-7). Specific wavelengths for use in quantitation are listed in Table 1. Other wavelengths may be used if they can provide the needed sensitivity and are corrected for spectral interference.

  - The elements listed in Sec. 1.1 must be calibrated. Flush the system for at least 60 seconds between each standard and sample using the rinse blank. At least three integrations averages are used for both calibration standard and sample analyses.

  - Calibrate the instrument with calibration standards (see sec.10)

  - For all analytes and determinations, the laboratory must analyze an ICV (Secs 10.7), a CCV (Secs. ‎10.8) and CCB (Secs.‎ ‎10.8.1 ) after every ten samples and at the end of the analysis batch run.

  - Analyze the samples, verify QC and record the results.

DATA ANALYSIS AND CALCULATIONS

  - If dilutions were performed, the appropriate factors must be applied to sample values. All results should be reported with up to two significant figures.

  - For samples with concentration above 500 ppm Na/Ca /Mg or/and above K 200 ppm the following samples must be re-analyzed due to the memory effect.

  - According to Appendix 1 to validation report (Nov 2024), due to spectral interferences:

    - K can be reported in calibration limits without dilution up to 1000ppm Na. Above this concertation, dilution is needed.

    - Sr can be reported in calibration limit without dilution up to 500ppm Ca. Above this concertation, dilution is needed.

WASTE MANAGEMENT

See "Laboratory Safety Working Procedure" W-03

REFERENCES

  - Inductively Coupled Plasma-Atomic Emission Spectrometry 6010D Revision 5. July 2018

  - W-12 "Instrument start up – shut down procedure"

  - Laboratory Safety Working Procedure" W-03

  - Validation report SOP 5 ICP-OES Aug 2023

  - Appendix 1 to validation report SOP 5 ICP-OES Nov 2024

  - Appendix 2 to validation report SOP 5 ICP-OES Dec 2025

TABLES, APPENDIX

TABLE 1. Recommended wavelengths

| Element | LLOQ [ppm] | Wavelength (nm)*1 | Report Range (mg/L) |
| --- | --- | --- | --- |
| Na | 1 | 589.592 588.995 589.592R 330.298R | 1-5 1-5 5-500 50-500 |
| K | 0.5 | 766.491 769.896 766.491R 769.896R | 0.5-5 0.5-5 5-500 5-500 |
| Element | LLOQ [ppm] | Wavelength (nm)*1 | Report Range (mg/L) |
| Mg | 0.1 | 285.213 279.079 279.079R 285.213R | 0.1-50 0.1-50 50-500 50-500 |
| Ca | 0.5 | 393.366 396.847 317.933 422.673 | 0.5-1 0.5-5 1-500 10-200 |
| Al | 0.05 | 167.078 396.152 | 0.05-1 0.05-10 |
| As | 0.05 | 189.042 193.759 | 0.05-5 0.05-10 |
| B | 0.05 | 249.773 249.677 182.641 | 0.05-10 0.05-10 0.05-5 |
| Ba | 0.01 | 233.527 455.404 | 0.01-10 0.01-10 |
| Be | 0.01 | 234.861 313.042 313.107 | 0.01-10 0.01-1 0.01-10 |
| Cd | 0.01 | 214.438 226.502 | 0.01-1 0.01-10 |
| Cr | 0.05 | 267.716 283.563 | 0.05-5 1-10 |
| Co | 0.01 | 228.616 238.892 | 0.01-10 0.01-5 |
| Cu | 0.05 | 324.754 224.700 327.396 | 0.05-10 0.05-1 0.05-1 |
| Fe | 0.05 | 238.204 259.941 | 0.05-1 0.05-10 |
| Li | 0.01 | 670.780R 670.780 | 0.01-10 0.01-10 |
| Mn | 0.05 | 257.611 259.373 | 0.05-10 0.05-5 |
| Ni | 0.05 | 221.648 231.604 232.003 | 0.5-1 0.5-10 0.5-10 |
| Pb | 0.05 | 213.857 283.305 | 0.05-1 0.5-10 |
| Se | 0.05 | 196.090 204.050 | 0.05-5 0.5-10 |
| Sr | 0.05 | 407.771 421.552 | 0.05-1 0.5-10 |
| Zn | 0.05 | 213.856 206.200 | 0.05-5 0.05-10 |

*1 R - radial view

*2 All curves must include blank

TABLE 2. Summary of Preparation Procedure of Standards and QC Samples

| Standard type | Concentration | Stock standard solution | Stock volume | Flask volume | HNO3 (65%) volume | IS (Y stock 20 mg/L) | Validity |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | mg/L |  | ml | ml | ml | ml |  |
| CAL | 200 | Na 1000 mg/L K 1000mg/L Mg 1000 mg/L Ca 1000 mg/L | 10 10 10 10 | 50 | 2.5 | 2.5 | One month |
| CAL | 100 | Na 1000 mg/L K 1000mg/L Mg 1000 mg/L Ca 1000 mg/L | 5 5 5 5 | 50 | 2.5 | 2.5 | One month |
| CAL CCV | 50 | Na 1000 mg/L K 1000mg/L Mg 1000 mg/L Ca 1000 mg/L | 2.5 2.5 2.5 2.5 | 50 | 2.5 | 2.5 | One month |
| CAL | 10 | ICP-VIII multi 100 mg/L As 1000 mg/L | 5 0.5 | 50 | 2.5 | 2.5 | One month |
| CAL | 5 | ICP-VIII multi 100 mg/L As 1000 mg/L | 2.5 0.25 | 50 | 2.5 | 2.5 | One month |
| CAL CCV | 1 | ICP-VIII multi 100 mg/L As 1000 mg/L | 0.5 0.05 | 50 | 2.5 | 2.5 | One month |
| CAL | 0.5 | ICP-VIII multi 100 mg/L As 1000 mg/L | 0.25 0.025 | 50 | 2.5 | 2.5 | One month |
| CAL | 0.1 | 1 mg/L ICP-VIII Multi+As | 5 | 50 | 2.5 | 2.5 | One month |
| CAL | 0.05 | 1 mg/L ICP-VIII Multi+As | 2.5 | 50 | 2.5 | 2.5 | One month |
| CAL | 0.01 | 1 mg/L ICP-VIII Multi+As | 0.5 | 50 | 2.5 | 2.5 | One month |
| IS | 20 | Y 1000 mg/L | 2 | 100 |  |  | One month |
| CAL CCB |  | DIW | 50 | 50 | 2.5 | 2.5 |  |
| ICV Major |  | ICP multi-element standard stock 10 | 10 | 10 | 0.5 | 0.5 |  |
| ICV Traces |  | WW LFS solution #1 | 0.5 | 50 | 2.5 | 2.5 | One month |
| ICV Major |  | WW LFS solution #1 | 5 | 50 | 2.5 | 2.5 | One month |
| Blank Rinse |  | DIW | 1000 | 1000 | 50 |  |  |
| Method Blank |  | DIW | 10 | 10 | 0.5 | 0.5 |  |

APPENDIX 1. Example for Sequence Formation

| # ICB # 0.01 mg/L # 0.05 mg/L # 0.1 mg/L # 0.5 mg/L # 1.0 mg/L # 5.0 mg/L # 10 mg/L # 50 mg/L # 100 mg/L # 200 mg/L | # Calibration Curve |
| --- | --- |
| # CCB # CCV # ICV |  |
| # 10 samples # Method blank |  |
| # CCB # CCV # ICV |  |
| # 10 samples |  |
| # CCB # CCV |  |